## 020627
YRDAY | SEQ | TYPE | Filenames for 02178 (June 27 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Hampton
|  |  | Optical Bench Temp: | not recorded
564449 - Collected IRFPA Image Table
02178 | AA | LVL | Level
02178 | AA | LVL | I020627_095316_090_B.FIT | 6 | 500 | I | [5 empty] | DAC XB8
02178 | AA | LVL | I020627_095319_523_B.FIT | 6 | 500 | I | [5 empty] | FPA temp ~120K, Bench ~150k
02178 | AA | LVL | I020627_123323_640_B.FIT | 6 | 501 | I | [5 empty] | DAC XB8
02178 | AA | LVL | I020627_123327_106_B.FIT | 6 | 501 | I | [5 empty] | FPA temp ~115K, Bench ~145k
02178 | AA | LVL | I020627_135133_073_B.FIT | 6 | 502 | I
02178 | AA | LVL | I020627_135136_623_B.FIT | 6 | 502 | I | [5 empty] | FPA temp ~113K (4 images)
02178 | AB | MSK | Focus
02178 | AB | MSK | I020627_135510_390_B.FIT | 4 | 503 | I | [5 empty] | Bias
02178 | AB | MSK | I020627_135513_240_B.FIT | 4 | 503 | I
02178 | AB | MSK | I020627_135516_123_B.FIT | 4 | 503 | I
02178 | AB | MSK | I020627_135520_273_B.FIT | 4 | 503 | I
02178 | AB | MSK | I020627_135613_890_B.FIT | 4 | 504 | I | [5 empty] | FIRST LIGHT!
02178 | AB | MSK | I020627_135616_740_B.FIT | 4 | 504 | I
02178 | AB | MSK | I020627_135619_623_B.FIT | 4 | 504 | I
02178 | AB | MSK | I020627_135623_723_B.FIT | 4 | 504 | I
02178 | AB | MSK | I020627_140039_490_B.FIT | 4 | 505 | I | [5 empty] | Slightly better focus
02178 | AB | MSK | I020627_140042_340_B.FIT | 4 | 505 | I
02178 | AB | MSK | I020627_140045_223_B.FIT | 4 | 505 | I
02178 | AB | MSK | I020627_140049_423_B.FIT | 4 | 505 | I
02178 | AB | MSK | I020627_140506_406_B.FIT | 4 | 506 | I | [5 empty] | Cover bottom pinhole; Got bottom pinhole!
02178 | AB | MSK | I020627_140509_256_B.FIT | 4 | 506 | I
02178 | AB | MSK | I020627_140512_140_B.FIT | 4 | 506 | I
02178 | AB | MSK | I020627_140516_240_B.FIT | 4 | 506 | I
02178 | AB | MSK | I020627_140941_090_B.FIT | 4 | 507 | I | [5 empty] | All, 12.0A
02178 | AB | MSK | I020627_140943_923_B.FIT | 4 | 507 | I
02178 | AB | MSK | I020627_140946_823_B.FIT | 4 | 507 | I
02178 | AB | MSK | I020627_140950_623_B.FIT | 4 | 507 | I
02178 | AB | MSK | I020627_141333_756_B.FIT | 4 | 508 | I | [5 empty] | New bias frame, 12.0A
02178 | AB | MSK | I020627_141336_640_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_141339_473_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_141343_540_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_141752_956_B.FIT | 4 | 509 | I | [5 empty] | Moved block, 12.0A
02178 | AB | MSK | I020627_141755_840_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_141758_673_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_141802_740_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142610_006_B.FIT | 4 | 510 | I | [5 empty] | 50 micron pinhole, ~ 60 DN, 15.0A
02178 | AB | MSK | I020627_142612_856_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142615_740_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142619_640_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142931_706_B.FIT | 4 | 511 | I | [5 empty] | Moved focus, 15.0A
02178 | AB | MSK | I020627_142934_556_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142937_440_B.FIT | 4 |  | I
02178 | AB | MSK | I020627_142941_340_B.FIT | 4 |  | I
02178 | AC | MSK | Focus
02178 | AC | MSK | I020627_143849_506_B.FIT | 4 | 512 | I | [5 empty] | START X-SLIT SCAN; Bias
02178 | AC | MSK | I020627_143852_340_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_143855_223_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_143859_173_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_143948_656_B.FIT | 4 | 513 | I |  | 0 |  |  |  | 500 micron, 12.0A
02178 | AC | MSK | I020627_143951_490_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_143954_373_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_143958_373_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144509_856_B.FIT | 4 | 514 | I |  | 100 |  |  |  | 500 micron, 12.0A, ~3700 DN
02178 | AC | MSK | I020627_144512_740_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144515_573_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144519_640_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144743_006_B.FIT | 4 | 515 | I |  | 200
02178 | AC | MSK | I020627_144745_840_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144748_740_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144752_540_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144859_006_B.FIT | 4 | 516 | I |  | 300
02178 | AC | MSK | I020627_144901_840_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144904_723_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144908_590_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144953_090_B.FIT | 4 | 517 | I |  | 400
02178 | AC | MSK | I020627_144955_923_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_144958_806_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145002_623_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145056_256_B.FIT | 4 | 518 | I |  | 500
02178 | AC | MSK | I020627_145059_090_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145101_973_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145105_990_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145242_806_B.FIT | 4 | 519 | I |  | 0
02178 | AC | MSK | I020627_145245_656_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145248_540_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145252_440_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145412_756_B.FIT | 4 | 520 | I |  | -100
02178 | AC | MSK | I020627_145415_590_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145418_490_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145422_490_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145507_840_B.FIT | 4 | 521 | I |  | -200
02178 | AC | MSK | I020627_145510_673_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145513_556_B.FIT | 4 |  | I
02178 | AC | MSK | I020627_145517_656_B.FIT | 4 |  | I
02178 | AD | MSK | Focus
02178 | AD | MSK | I020627_145650_006_B.FIT | 4 | 522 | I |  | -300 |  |  |  | Remove pinholes
02178 | AD | MSK | I020627_145652_840_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_145655_740_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_145659_740_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150016_806_B.FIT | 4 | 523 | I |  | n/a |  |  |  | Pinhole out, slit 1 millimeter
02178 | AD | MSK | I020627_150019_656_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150022_540_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150026_490_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150614_390_B.FIT | 4 | 524 | I |  | 4.6 |  |  |  | Slit 0.5 millimeter, moved by geometric view
02178 | AD | MSK | I020627_150617_240_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150620_123_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_150624_023_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_151058_023_B.FIT | 4 | 525 | I | [5 empty] | Rotate slit 90 degrees
02178 | AD | MSK | I020627_151100_856_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_151103_756_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_151107_706_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_151955_073_B.FIT | 4 | 526 | I | 17 |  |  |  |  | Moved focus 2 millimeters away from slit
02178 | AD | MSK | I020627_151957_923_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_152000_806_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_152004_706_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_152143_690_B.FIT | 4 | 527 | I | -30 |  |  |  |  | Not much change
02178 | AD | MSK | I020627_152146_540_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_152149_423_B.FIT | 4 |  | I
02178 | AD | MSK | I020627_152153_240_B.FIT | 4 |  | I
02178 | AE | MSK | Focus
02178 | AE | MSK | I020627_152659_490_B.FIT | 4 | 528 | I | [5 empty] | Add pinholes; Slit to 100 micron, still 90 degree rotation
02178 | AE | MSK | I020627_152702_323_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_152705_206_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_152709_256_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_152950_840_B.FIT | 4 | 529 | I |  | -2 |  |  |  | -2 millimeter horizontal (to the right)
02178 | AE | MSK | I020627_152953_723_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_152956_556_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153000_756_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153239_023_B.FIT | 4 | 530 | I |  | -4
02178 | AE | MSK | I020627_153241_856_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153244_740_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153248_690_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153351_640_B.FIT | 4 | 531 | I |  | -6
02178 | AE | MSK | I020627_153354_523_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153357_356_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153401_506_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153535_190_B.FIT | 4 | 532 | I |  | -8
02178 | AE | MSK | I020627_153538_023_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153540_923_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153544_923_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153710_740_B.FIT | 4 | 533 | I |  | -10
02178 | AE | MSK | I020627_153713_623_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153716_456_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153720_523_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153904_190_B.FIT | 4 | 534 | I |  | -12
02178 | AE | MSK | I020627_153907_040_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153909_923_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_153913_923_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_154050_490_B.FIT | 4 | 535 | I |  | 2
02178 | AE | MSK | I020627_154053_340_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_154056_223_B.FIT | 4 |  | I
02178 | AE | MSK | I020627_154100_173_B.FIT | 4 |  | I
02178 | AF | MSK | Focus
02178 | AF | MSK | I020627_154238_540_B.FIT | 4 | 536 | I | 0 | 4 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154241_390_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154244_273_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154248_173_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154332_840_B.FIT | 4 | 537 | I | 0 | 6 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154335_690_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154338_573_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154342_473_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154459_123_B.FIT | 4 | 538 | I | 0 | 8 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154501_973_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154504_856_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154508_556_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154610_890_B.FIT | 4 | 539 | I | 0 | 10 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154613_740_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154616_623_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154620_523_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154717_290_B.FIT | 4 | 540 | I | 0 | 12 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154720_140_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154723_023_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154726_923_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154841_640_B.FIT | 4 | 541 | I | 0 | 14 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154844_473_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154847_373_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154851_323_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154933_340_B.FIT | 4 | 542 | I | 0 | 16 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_154936_190_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154939_073_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_154942_973_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155029_873_B.FIT | 4 | 543 | I | 0 | 18 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_155032_723_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155035_606_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155039_606_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155226_906_B.FIT | 4 | 544 | I | 0 | 20 | 0 |  |  | Searching for image.
02178 | AF | MSK | I020627_155229_756_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155232_640_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_155236_340_B.FIT | 4 |  | I
02178 | AF | MSK | I020627_160247_740_B.FIT | 4 | 545 | I | 0 | 0.12 | 0 |  |  | Searching for image.
02178 | AG | MSK | Focus
02178 | AG | MSK | I020627_160250_590_B.FIT | 4 |  | I | [5 empty] | Saw image and moved table by hand.
02178 | AG | MSK | I020627_160253_473_B.FIT | 4 |  | I
02178 | AG | MSK | I020627_160257_523_B.FIT | 4 |  | I
02178 | AH | MSK | Focus
02178 | AH | MSK | I020627_161032_340_B.FIT | 4 | 546 | I | 0 | 0.12 | 0 |  |  | Moved again by hand.
02178 | AH | MSK | I020627_161035_173_B.FIT | 4 |  | I
02178 | AH | MSK | I020627_161038_073_B.FIT | 4 |  | I
02178 | AH | MSK | I020627_161041_823_B.FIT | 4 |  | I
02178 | AI | MSK | Focus
02178 | AI | MSK | I020627_161220_090_B.FIT | 4 | 547 | I | 0 | 0 | 0 |  |  | RESET TO ZERO. Rotated slit to vertical.
02178 | AI | MSK | I020627_161222_940_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161225_823_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161229_973_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161555_390_B.FIT | 4 | 548 | I | 0 | 0.5 | 0 |  |  | Moved X-Slit 500 micron "south".
02178 | AI | MSK | I020627_161558_240_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161601_123_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161605_223_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161946_723_B.FIT | 4 | 549 | I | 0 | -0.5 | 0 |  |  | Moved X-Slit 500 micron "north".
02178 | AI | MSK | I020627_161949_573_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161952_456_B.FIT | 4 |  | I
02178 | AI | MSK | I020627_161956_506_B.FIT | 4 |  | I
02178 | AJ | MSK | Focus
02178 | AJ | MSK | I020627_163024_973_B.FIT | 4 | 550 | I | 0 | 0 | 0 |  |  | Slit vertical, manually adjusted table, 1mm slit.
02178 | AJ | MSK | I020627_163027_806_B.FIT | 4 |  | I
02178 | AJ | MSK | I020627_163030_706_B.FIT | 4 |  | I
02178 | AJ | MSK | I020627_163034_706_B.FIT | 4 |  | I
02178 | AK | MSK | Focus
02178 | AK | MSK | I020627_163218_540_B.FIT | 4 | 551 | I | 0 | 0 | 0 |  |  | 100 micron slit.
02178 | AK | MSK | I020627_163221_390_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163224_273_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163228_273_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163401_340_B.FIT | 4 | 552 | I | 0 | 0.1 | 0 |  |  | Moved X-Slit 100 micron "south".
02178 | AK | MSK | I020627_163404_173_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163407_073_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163411_073_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163740_890_B.FIT | 4 | 553 | I | 0 | -0.1 | 0 |  |  | Moved X-Slit 100 micron "north".
02178 | AK | MSK | I020627_163743_740_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163746_623_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163750_523_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163946_123_B.FIT | 4 | 554 | I | 0 | -0.2 | 0 |  |  | Moved X-Slit 200 micron "north".
02178 | AK | MSK | I020627_163948_956_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163951_856_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_163955_656_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_164157_290_B.FIT | 4 | 555 | I | 0 | -0.3 | 0 |  |  | Moved X-Slit 300 micron "north".
02178 | AK | MSK | I020627_164200_173_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_164203_006_B.FIT | 4 |  | I
02178 | AK | MSK | I020627_164207_256_B.FIT | 4 |  | I
02178 | AL | MSK | Focus
02178 | AL | MSK | I020627_164840_306_B.FIT | 4 | 556 | I | 0 | -0.2 | 0 |  |  | Zeroed at 200 micron "north". FPA at 108K.
02178 | AL | MSK | I020627_164843_156_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_164846_040_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_164850_240_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165033_273_B.FIT | 4 | 557 | I | 0 | -0.3 | 0 |  |  | Moved X-Slit 300 micron "north".
02178 | AL | MSK | I020627_165036_123_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165039_006_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165042_956_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165256_223_B.FIT | 4 | 558 | I | 0 | -0.4 | 0 |  |  | Moved X-Slit 400 micron "north".
02178 | AL | MSK | I020627_165259_073_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165301_956_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165305_906_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165444_723_B.FIT | 4 | 559 | I | 0 | -0.5 | 0 |  |  | Moved X-Slit 500 micron "north".
02178 | AL | MSK | I020627_165447_573_B.FIT | 4 |  | I | [5 empty] | Set to 300 micron "north" and zero at highest value.
02178 | AL | MSK | I020627_165450_456_B.FIT | 4 |  | I
02178 | AL | MSK | I020627_165454_556_B.FIT | 4 |  | I
02178 | AM | MSK | Focus
02178 | AM | MSK | I020627_170505_106_B.FIT | 4 | 560 | I | 0 | 0 | 0 |  |  | 500 micron pixels
02178 | AM | MSK | I020627_170507_940_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_170510_840_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_170514_640_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171153_106_B.FIT | 4 | 561 | I | -12 | 0 | 0 |  |  | Moved Z-focus -12mm towards chamber.
02178 | AM | MSK | I020627_171155_940_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171158_840_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171202_590_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171544_873_B.FIT | 4 | 562 | I | -30 | 0 | 0
02178 | AM | MSK | I020627_171547_756_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171550_590_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171554_656_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171810_840_B.FIT | 4 | 563 | I | -50 | 0 | 0
02178 | AM | MSK | I020627_171813_673_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171816_573_B.FIT | 4 |  | I
02178 | AM | MSK | I020627_171820_373_B.FIT | 4 |  | I
02178 | AN | MSK | Focus
02178 | AN | MSK | I020627_172736_373_B.FIT | 4 | 564 | I | -60 | 0 | 0 |  |  | Aligned and changed source to 15.0 Amp.
02178 | AN | MSK | I020627_172739_206_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_172742_090_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_172746_190_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_173529_723_B.FIT | 4 | 565 | I | -100 | 0 | 0
02178 | AN | MSK | I020627_173532_556_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_173535_456_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_173539_206_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_174021_640_B.FIT | 4 | 566 | I |  | 0 | 0 |  |  | Cover bottom pinhole (looks like bottom signal dissappeared).
02178 | AN | MSK | I020627_174024_490_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_174027_373_B.FIT | 4 |  | I
02178 | AN | MSK | I020627_174031_223_B.FIT | 4 |  | I
02178 | AO | MSK | Focus
02178 | AO | MSK | I020627_175314_573_B.FIT | 4 | 567 | I |  | 0 | 0 |  |  | HeNe Laser on.
02178 | AO | MSK | I020627_175317_423_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_175320_306_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_175324_306_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_175724_723_B.FIT | 4 | 568 | I |  | 0 | 0 |  |  | Manually shift 500 micron. Source (lamp) = 12.0 Amp. T (def) = 106.5K.
02178 | AO | MSK | I020627_175727_573_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_175730_456_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_175734_306_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180209_573_B.FIT | 4 | 569 | I |  | 0 | 0 |  |  | Slit vertical, 100 micron.
02178 | AO | MSK | I020627_180212_423_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180215_306_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180219_406_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180510_406_B.FIT | 4 | 570 | I |  | 0 | 0 |  |  | 500 micron pinholes.
02178 | AO | MSK | I020627_180513_240_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180516_140_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180519_940_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180923_473_B.FIT | 4 | 571 | I |  | 0 | 0 |  |  | Refilled the aperature.
02178 | AO | MSK | I020627_180926_306_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180929_206_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_180933_206_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181114_406_B.FIT | 4 | 572 | I |  | 0 | 0 |  |  | More refilling.
02178 | AO | MSK | I020627_181117_240_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181120_140_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181123_990_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181757_756_B.FIT | 4 | 573 | I |  | 0 | 0 |  |  | Tilted table.
02178 | AO | MSK | I020627_181800_606_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181803_490_B.FIT | 4 |  | I
02178 | AO | MSK | I020627_181807_490_B.FIT | 4 |  | I
02178 | AP | MSK | Focus
02178 | AP | MSK | I020627_181932_623_B.FIT | 4 | 574 | I |  | 0 | 0 |  |  | Manually adjusted table.
02178 | AP | MSK | I020627_181935_456_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_181938_356_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_181942_356_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182450_090_B.FIT | 4 | 575 | I |  | 0 | 0 |  |  | 100 micron slit only. Source = 12.0 Amp.
02178 | AP | MSK | I020627_182452_940_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182455_823_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182500_023_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182903_906_B.FIT | 4 | 576 | I |  | 0 | 0 |  |  | Moved + X-axis 300 micron.
02178 | AP | MSK | I020627_182906_790_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182909_623_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_182913_823_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183045_156_B.FIT | 4 | 577 | I |  | 0 | 0 |  |  | Moved - X-axis -250 micron (del = 550 micron).
02178 | AP | MSK | I020627_183047_990_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183050_890_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183054_890_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183134_406_B.FIT | 4 | 578 | I |  | 0 | 0 |  |  | to -56 micron.
02178 | AP | MSK | I020627_183137_240_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183140_123_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183144_190_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183417_706_B.FIT | 4 | 579 | I |  | 0.1 | 0
02178 | AP | MSK | I020627_183420_540_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183423_440_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183427_440_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183548_040_B.FIT | 4 | 580 | I |  | 0.2 | 0
02178 | AP | MSK | I020627_183550_873_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183553_773_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183557_523_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183643_456_B.FIT | 4 | 581 | I |  | 0.3 | 0
02178 | AP | MSK | I020627_183646_340_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183649_173_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183653_240_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183740_256_B.FIT | 4 | 582 | I |  | 0.4 | 0
02178 | AP | MSK | I020627_183743_140_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183745_973_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183750_173_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183828_190_B.FIT | 4 | 583 | I |  | 0.5 | 0
02178 | AP | MSK | I020627_183831_023_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183833_923_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183837_673_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183917_390_B.FIT | 4 | 584 | I |  | 0.6 | 0
02178 | AP | MSK | I020627_183920_223_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183923_106_B.FIT | 4 |  | I
02178 | AP | MSK | I020627_183927_156_B.FIT | 4 |  | I
02178 | AQ | MSK | Focus
02178 | AQ | MSK | I020627_184345_206_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotate slit 90 degrees. 500 micron slit.
02178 | AQ | MSK | I020627_184646_440_B.FIT | 1 | n/a | I |  |  | 0 |  |  | 500 micron slit.
02178 | AQ | MSK | I020627_184944_390_B.FIT | 1 | n/a | I |  |  | 0 |  |  | 100 micron slit.
02178 | AQ | MSK | I020627_185249_090_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Blocked image around slit/center (106 y).
02178 | AQ | MSK | I020627_185616_256_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotate slit to 45 degrees.
02178 | AQ | MSK | I020627_185729_040_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotate slit to 90 degrees.
02178 | AQ | MSK | I020627_190058_540_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Reset slit.
02178 | AQ | MSK | I020627_190202_540_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotate slit 45 degrees.
02178 | AQ | MSK | I020627_190317_340_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Change slit to 2mm.
02178 | AQ | MSK | I020627_190433_356_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotate closer to vertical.
02178 | AQ | MSK | I020627_190709_056_B.FIT | 1 | n/a | I |  |  | 0 |  |  | +2mm south
02178 | AQ | MSK | I020627_190810_156_B.FIT | 1 | n/a | I |  |  | 0 |  |  | +10 degrees.
02178 | AQ | MSK | I020627_190908_890_B.FIT | 1 | n/a | I |  |  | 0 |  |  | +10 degrees.
02178 | AQ | MSK | I020627_191010_806_B.FIT | 1 | n/a | I |  |  | 0 |  |  | +100 micron south. Same angle.
02178 | AQ | MSK | I020627_191119_373_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Rotated to vertical slit.
02178 | AQ | MSK | I020627_191627_740_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Changed to 500 micron pinholes.
02178 | AQ | MSK | I020627_191844_040_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Changed focus by - 30 mm.
02178 | AQ | MSK | I020627_192033_590_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Changed foccus to -30mm.
02178 | AQ | MSK | I020627_192405_990_B.FIT | 1 | n/a | I |  |  | 0 |  |  | Blocked both pinholes.
02178 | AQ | MSK | I020627_193424_023_B.FIT | 1 | n/a | I |  | 0.5 | 0 |  |  | X-slit at 500 micron. Move Klinger vertical stage.
02178 | AQ | MSK | I020627_193600_473_B.FIT | 1 | n/a | I |  | 0.5 | 0 |  |  | Down to 100 micron slit.
02178 | AQ | MSK | I020627_193749_223_B.FIT | 1 | n/a | I |  | 0.5 | 0.2 |  |  | Move down to 2mm.
02178 | AQ | MSK | I020627_194122_440_B.FIT | 1 | n/a | I |  | 0.5 | 12 |  |  | Move down to 12mm.
02178 | AQ | MSK | I020627_194343_806_B.FIT | 1 | 600 | I |  | 0.5 | -12
02178 | AQ | MSK | I020627_194447_756_B.FIT | 1 | 601 | I |  | 0.5 | -24
02178 | AQ | MSK | I020627_194551_006_B.FIT | 1 | 602 | I |  | 0.5 | -40
02178 | AQ | MSK | I020627_195122_023_B.FIT | 1 | 603 | I |  | 0.5 | -60 |  |  | No image.
02178 | AQ | MSK | I020627_195331_623_B.FIT | 1 | 604 | I |  | 0.5 | 0 |  |  | Moved table back. 100 micron slit.

## 020628
YRDAY | SEQ | TYPE | Filenames for 02179 (June 28 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Van Cleve
|  |  | Optical Bench Temp: | not recorded
564449 - Collected IRFPA Image Table
02179 | AA | LVL | Level
02179 | AA | LVL | I020628_095824_173_B.FIT | 6 |  | I | [5 empty] | Dark. Reference = wood box in front of tungsten lamp.
02179 | AA | LVL | I020628_095828_606_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | Allmode darks
02179 | AB | DRKW | I020628_100219_373_B.FIT | 1 | 600 | I | [5 empty] | Dark. Temp (FPA) = ~92K, Temp (Optical Bench) = ~125K
02179 | AB | DRKW | I020628_100222_156_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100225_106_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100227_890_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100230_840_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100233_623_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100236_573_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100240_506_B.FIT | 1 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100407_373_B.FIT | 2 | 601 | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100408_856_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100410_340_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100411_640_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100413_106_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100414_590_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100416_073_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100418_723_B.FIT | 2 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100543_523_B.FIT | 3 | 602 | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100544_273_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100545_006_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100545_756_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100546_490_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100547_223_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100547_940_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100549_940_B.FIT | 3 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100719_273_B.FIT | 4 | 603 | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100722_156_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100724_990_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100729_190_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100839_790_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100842_640_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100845_523_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_100849_673_B.FIT | 4 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101018_590_B.FIT | 5 | 604 | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101021_540_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101024_356_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101027_323_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101030_090_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101033_056_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101035_823_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101038_740_B.FIT | 5 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101207_023_B.FIT | 6 | 605 | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101210_540_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101214_123_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101217_590_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101221_090_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101224_573_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101228_056_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AB | DRKW | I020628_101231_406_B.FIT | 6 |  | I | [5 empty] | Dark
02179 | AC | CHK | Memory Check
02179 | AC | CHK | I020628_101654_106_B.FIT | 7 | 1 | I | [5 empty] | Memory check mode - No data.
02179 | AC | CHK | I020628_101657_023_B.FIT | 7 |  | I | [5 empty] | Memory check mode - No data.
02179 | AD | MSK | Focus
02179 | AD | MSK | I020628_102610_190_B.FIT | 1 | 606 | I |  | 0 | 0 |  |  | Focus frame; Oops - focus at some low level: Retry
02179 | AD | MSK | I020628_102614_340_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_102803_156_B.FIT | 1 | 607 | I | 0 | 0 | 0 |  |  | Moved focus;
02179 | AD | MSK | I020628_102807_290_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_103026_056_B.FIT | 1 | 608 | I | +50 | 0 | 0 |  |  | Moved focus; Looks more out of focus: Separation ~30 pixels
02179 | AD | MSK | I020628_103030_190_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_103336_940_B.FIT | 1 | 609 | I | -50 |  |  |  |  | Moved focus; Separation ~23 pixels; Getting better but still way out of focus!
02179 | AD | MSK | I020628_103341_140_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_103552_856_B.FIT | 1 | 610 | I | -50 |  |  |  |  | Operator error on control procedure (pressed <CR> at mode prompt). Quit procedure and restart.
02179 | AD | MSK | I020628_103557_006_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_103737_140_B.FIT | 1 | 611 | I | -50
02179 | AD | MSK | I020628_103741_323_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_104122_090_B.FIT | 1 | 612(?) | I | n/a | n/a | n/a |  |  | Jostle table; No info on X-slit and Vert; Restarting Image Ops (enter CR)
02179 | AD | MSK | I020628_104126_290_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_104615_156_B.FIT | 1 | 620 | I | n/a | n/a | n/a |  |  | Exposure ID (1,0) = (2, 108)
02179 | AD | MSK | I020628_104619_340_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_105146_940_B.FIT | 1 | 621 | I | 0 | 0 | 0 |  |  | Back off 30mm from tank source. Source = 11Amp.
02179 | AD | MSK | I020628_105151_073_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_105642_640_B.FIT | 1 | 622 | I | [5 empty] | Move in towards tank source. Got a very narrow and dim line at a different location.
02179 | AD | MSK | I020628_105646_823_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_105814_590_B.FIT | 1 | 623 | I | [5 empty] | Open slit at same position; Line a little wider (~4 pixels) but brighter
02179 | AD | MSK | I020628_105818_773_B.FIT | 1 |  | I
02179 | AD | MSK | I020628_110323_723_B.FIT | 1 | 624 | I | [5 empty] | Adjusted table. Slit at 200 micron. Very saturated image. Source = 1Amp.
02179 | AD | MSK | I020628_110514_390_B.FIT | 1 | 625 | I | [5 empty] | Slit at 100 micron.
02179 | AD | MSK | I020628_110746_956_B.FIT | 1 | 626 | I | [5 empty] | Source = 11Amp. Still saturated. About 7mm out of focus.
02179 | AD | MSK | I020628_111056_156_B.FIT | 1 | 627 | I | 7
02179 | AD | MSK | I020628_111215_040_B.FIT | 1 | 628 | I | 9 |  |  |  |  | Down to ~3 pixels in image.
02179 | AE | MSK | Focus
02179 | AE | MSK | I020628_111349_790_B.FIT | 4 | 629 | I | 9 |  |  |  |  | Close slit to 50mm. Change mode for full resolution. FWHM (col n/a) = 7.5 pix
02179 | AE | MSK | I020628_111352_623_B.FIT | 4
02179 | AE | MSK | I020628_111355_523_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111359_573_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111651_990_B.FIT | 4 | 630 | I | 11 |  |  |  |  | FWHM (col 425) = 6.0 pix = improvement
02179 | AE | MSK | I020628_111654_840_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111657_723_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111701_623_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111920_590_B.FIT | 4 | 631 | I | 13 |  |  |  |  | FWHM (col 425) = 5.08 pix = improvement
02179 | AE | MSK | I020628_111923_440_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111926_323_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_111930_423_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_112043_506_B.FIT | 4 | 632 | I | 15 |  |  |  |  | FWHM (col 425) = 4.83; Best focus and found source; Slit jaws at 50 micron.
02179 | AE | MSK | I020628_112046_356_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_112049_240_B.FIT | 4 |  | I
02179 | AE | MSK | I020628_112053_090_B.FIT | 4 |  | I
02179 | AF | MSK | Focus
02179 | AF | MSK | I020628_112643_906_B.FIT | 4 | 633 | I | [5 empty] | Begin Aligment procedure. Open slit jaws to 100 micron and rotate 20 degree from horizontal.
02179 | AF | MSK | I020628_112646_756_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_112649_640_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_112653_490_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113005_856_B.FIT | 4 | 634 | I | 15 | 2 | n/a |  |  | Translate 3mm; Lost some signal and focus but about back to near vertical.
02179 | AF | MSK | I020628_113008_706_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113011_590_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113015_490_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113216_240_B.FIT | 4 | 635 | I | 15 | -2 | n/a |  |  | Translate 2mm the other way.
02179 | AF | MSK | I020628_113219_090_B.FIT | 4
02179 | AF | MSK | I020628_113221_973_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113226_173_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113723_640_B.FIT | 4 | 636 | I | 15 | -3 | n/a |  |  | Rotate slit to 45 degrees from horizontal.
02179 | AF | MSK | I020628_113726_490_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113729_373_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113733_473_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113839_723_B.FIT | 4 | 637 | I | 15 | -3 | n/a |  |  | Rotate slit to 60 degrees from horizontal. Slit still at 100 microns. Oops - we were not at 60 degrees.
02179 | AF | MSK | I020628_113842_556_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113845_440_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_113849_256_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114037_890_B.FIT | 4 | 638 | I | 15 | -2.5 | n/a |  |  | Rotate slit to 60 degrees from horizontal. Source = 11Amp.
02179 | AF | MSK | I020628_114040_740_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114043_623_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114047_723_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114226_940_B.FIT | 4 | 639 | I | 15 | -2.5 | n/a |  |  | Rotate slit to 75 degrees from horizontal.
02179 | AF | MSK | I020628_114229_790_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114232_673_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114236_673_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114400_090_B.FIT | 4 | 640 | I | 15 | -2.5 | n/a |  |  | Change slit to 50 microns.
02179 | AF | MSK | I020628_114402_940_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114405_823_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114409_923_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114709_090_B.FIT | 4 | 641 | I | 15 | -2.3 | n/a |  |  | Rotate slit to 90 degrees from horizontal. Keep slit at 50 microns.
02179 | AF | MSK | I020628_114711_940_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114714_823_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114718_923_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114851_090_B.FIT | 4 | 642 | I | 15 | -2.4 | n/a |  |  | Slit angle still 90 degrees. Image looks worse.
02179 | AF | MSK | I020628_114853_923_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114856_823_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114900_623_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114947_990_B.FIT | 4 | 643 | I | 15 | -2.2 | n/a |  |  | Looks good.
02179 | AF | MSK | I020628_114950_840_B.FIT | 4 |  | I
02179 | AF | MSK | I020628_114953_723_B.FIT | 4
02179 | AF | MSK | I020628_114957_723_B.FIT | 4 |  | I
02179 | AG | MSK | Focus
02179 | AG | MSK | I020628_115714_390_B.FIT | 4 | 644 | I | 15 | -2.2 | n/a |  |  | Swith to 3, 500-micron pinholes at vertical. Bottom 2 pinholes fully illuminated.
02179 | AG | MSK | I020628_115717_240_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_115720_123_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_115724_073_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_121106_806_B.FIT | 4 | 645 | I | 15 | -2.2 | n/a |  |  | Put in laser to look at alignment of boresight and adjusted table angle. Need another table adjustment.
02179 | AG | MSK | I020628_121109_656_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_121112_540_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_121116_440_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_121400_456_B.FIT | 4 | 646 | I | 15 | -2.2 | n/a |  |  | Slightly rotated about center pixel but no change to table angle. Lost some signal through all pinholes.
02179 | AG | MSK | I020628_121403_306_B.FIT | 4 |  | I | [5 empty] | Still have boresight and registration errors. Not ready for focus yet. Need to drop table to raise image.
02179 | AG | MSK | I020628_121406_190_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_121410_090_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_123531_206_B.FIT | 4 | 647 | I | 15 | -2.2 | n/a |  |  | Lowered table while maintaining boresight.
02179 | AG | MSK | I020628_123534_040_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_123536_940_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_123540_940_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_124516_790_B.FIT | 4 | 648 | I | 15 | -2.2 | n/a |  |  | Raise table (telescope). The spatial boresight is correct. Laser agrees with what image shows.
02179 | AG | MSK | I020628_124519_623_B.FIT | 4 |  | I | [5 empty] | A little off in spectral boresight (requires slight rotation of table). Upper pinhole barely illuminated.
02179 | AG | MSK | I020628_124522_506_B.FIT | 4 |  | I
02179 | AG | MSK | I020628_124526_456_B.FIT | 4 |  | I
02179 | AH | MSK | Focus
02179 | AH | MSK | I020628_135147_656_B.FIT | 6 | 649 | I | 15 | n/a | n/a |  |  | IRFPA turned off for 1/2-hour lunch break. Standard dark (diagnostic mode pair). Left quad = 3621 DN, Right quad = 4581
02179 | AH | MSK | I020628_135152_090_B.FIT | 6 |  | I
02179 | AI | DRKW | Darks
02179 | AI | DRKW | I020628_135404_290_B.FIT | 4 | 650 | I | 15 | n/a | n/a |  |  | Dark
02179 | AI | DRKW | I020628_135407_140_B.FIT | 4 |  | I
02179 | AI | DRKW | I020628_135410_023_B.FIT | 4 |  | I
02179 | AI | DRKW | I020628_135413_923_B.FIT | 4 |  | I
02179 | AJ | MSK | Focus
02179 | AJ | MSK | I020628_135915_690_B.FIT | 4 | 651 | I | 15 | n/a | n/a |  |  | Take a picture to make sure setup didn't change. Moved about 5 pixels high since lunch.
02179 | AJ | MSK | I020628_135918_540_B.FIT | 4 |  | I | [5 empty] | This means we are about 8 pixels too high. See if the extra box on the table caused the shift.
02179 | AJ | MSK | I020628_135921_423_B.FIT | 4 |  | I | [5 empty] | Took the box off the table
02179 | AJ | MSK | I020628_135925_573_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_140602_206_B.FIT | 4 | 652 | I | 15 | n/a | n/a |  |  | We are back to 3 pixels. The box caused the shift. The line width grew by 0.5 pix (out of ~9 pix).
02179 | AJ | MSK | I020628_140605_056_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_140607_940_B.FIT | 4
02179 | AJ | MSK | I020628_140611_890_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_140908_173_B.FIT | 4 | 653 | I | 15 | n/a | n/a |  |  | Tape measure removed from table (we may need it later). No change in image.
02179 | AJ | MSK | I020628_140911_023_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_140913_906_B.FIT | 4
02179 | AJ | MSK | I020628_140918_006_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_141653_790_B.FIT | 4 | 654 | I | 15 | 0 | 0 |  |  | Adjusted table for spectral boresight. Image shifted down by ~10 pix.
02179 | AJ | MSK | I020628_141656_623_B.FIT | 4 |  | I | [5 empty] | Lowest spectrum is brightest. In previous image, 2 lowest spectrums had ~equal brightness.
02179 | AJ | MSK | I020628_141659_506_B.FIT | 4 |  | I | [5 empty] | Is pinhole rotated or is source intensity different?
02179 | AJ | MSK | I020628_141703_606_B.FIT | 4
02179 | AJ | MSK | I020628_143517_556_B.FIT | 4 | 655 | I | 15 | 0 | 0 |  |  | Realigned the concentrating optic for the source. Lower 2 spectrums have about the same brighteness.
02179 | AJ | MSK | I020628_143520_390_B.FIT | 4 |  | I | [5 empty] | About 12 pixels low.
02179 | AJ | MSK | I020628_143523_290_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_143527_090_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_143720_340_B.FIT | 4 | 656 | I | 15 | 0.1 | 0 |  |  | Moved X-slit +100 microns to get more uniform illumination of pinholes.
02179 | AJ | MSK | I020628_143723_190_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_143726_073_B.FIT | 4 |  | I
02179 | AJ | MSK | I020628_143730_123_B.FIT | 4 |  | I
02179 | AK | MSK | Focus
02179 | AK | MSK | I020628_144353_856_B.FIT | 4 | 657 | I | 15 | 0.1 | 0 |  |  | Switch to 50 micron pinholes. Source still at 12 Amp. Try to align pinholes.
02179 | AK | MSK | I020628_144356_740_B.FIT | 4 |  | I | [5 empty] | Spectrums dim and still not equal intensity. Not sure if it's the source, illumination, or pinholes.
02179 | AL | MSK | Focus
02179 | AL | MSK | I020628_144359_573_B.FIT | 4 | [7 empty] | Try knife-edge.
02179 | AL | MSK | I020628_144403_640_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_144803_723_B.FIT | 4 | 658 | I | 15 | 0.05 | 0 |  |  | Moved X-slit back 50 microns. Spectrums are dimmer.
02179 | AL | MSK | I020628_144806_573_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_144809_456_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_144813_473_B.FIT | 4
02179 | AL | MSK | I020628_145107_656_B.FIT | 4 | 659 | I | 15 | 0.15 | 0 |  |  | Moved X-slit to 150 microns. Signal has increased.
02179 | AL | MSK | I020628_145110_506_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145113_390_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145117_290_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145542_306_B.FIT | 4 | 660 | I | 15 | 0.1 | 0 |  |  | Signal about the same. Try 200 microns
02179 | AL | MSK | I020628_145545_140_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145548_023_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145552_073_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145916_173_B.FIT | 4 | 661 | I | 15 | 0.2 | 0 |  |  | Signal increased. Center most intense and top is increasing relative to lower 2 spectrums.
02179 | AL | MSK | I020628_145919_006_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145921_906_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_145925_706_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150115_190_B.FIT | 4 | 662 | I | 15 | 0.25 | 0 |  |  | Intensity is about the same as previous frame. Try one more X-slit step to verify.
02179 | AL | MSK | I020628_150118_073_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150120_906_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150125_156_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150303_790_B.FIT | 4 | 663 | I | 15 | 0.3 | 0 |  |  | Plot of X-Slit vs DN using pixel (,). Plot indicates best X-Slit is 225 microns.
02179 | AL | MSK | I020628_150306_673_B.FIT | 4 |  | I | [5 empty] | See Excel spreadsheet HRI_Xslit_vs_DN_020628.xls.
02179 | AL | MSK | I020628_150309_506_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150313_756_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150934_106_B.FIT | 4 | 664 | I | 15 | 0.225 | 0 |  |  | X-Slit = 225 microns looks good!
02179 | AL | MSK | I020628_150936_956_B.FIT | 4 |  | I | [5 empty] | Rotation looks good for Zinc Selenide prism.
02179 | AL | MSK | I020628_150939_840_B.FIT | 4 |  | I
02179 | AL | MSK | I020628_150943_740_B.FIT | 4 |  | I
02179 | AM | MSK | Focus
02179 | AM | MSK | I020628_153228_540_B.FIT | 4 | 665 | I | 15 | 0.225 | 0 |  |  | Begin boresight test (dark-subtracted photometry). Use a 40x40 box centered on (650,244).
02179 | AM | MSK | I020628_153231_390_B.FIT | 4 |  | I | [5 empty] | Using dark image I020628_135404_290_B.FIT. Source at 12Amp.
02179 | AM | MSK | I020628_153234_273_B.FIT | 4 |  | I | [5 empty] | DN mean for box via DIVe = 15.0675
02179 | AM | MSK | I020628_153238_423_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_153421_073_B.FIT | 4 | 666 | I | 15 | 0.225 | 0 |  |  | Begin aperature masking. 1/2 aperature vertical, left masked.
02179 | AM | MSK | I020628_153423_956_B.FIT | 4 |  | I | [5 empty] | DN mean for box via DIVe = 10.09125
02179 | AM | MSK | I020628_153426_790_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_153430_990_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_153546_240_B.FIT | 4 | 667 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature horizontal, bottom masked.
02179 | AM | MSK | I020628_153549_090_B.FIT | 4 |  | I | [5 empty] | DN mean for box via DIVe = 6.6525. Pupil not being illuminated uniformly?
02179 | AM | MSK | I020628_153551_973_B.FIT | 4 |  | I | [5 empty] | Confident that boresight is aligned with the laser. Will also analyze DN via more robust IDL routine.
02179 | AM | MSK | I020628_153556_073_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_155043_073_B.FIT | 4 | 668 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature vertical, right masked WITH PAPER INSTEAD OF MYLAR!
02179 | AM | MSK | I020628_155045_923_B.FIT | 4 |  | I | [5 empty] | DN mean for box via DIVe = 15.0075
02179 | AM | MSK | I020628_155048_806_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_155052_706_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_155547_656_B.FIT | 4 | 669 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature horizontal, top masked.
02179 | AM | MSK | I020628_155550_506_B.FIT | 4 |  | I | [5 empty] | DN mean for box via DIVe = 37.40375
02179 | AM | MSK | I020628_155553_390_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_155557_490_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_160032_490_B.FIT | 4 | 670 | I | 15 | 0.225 | 0 |  |  | Recollect bottom masked (Exp ID 667); it may be suspect.
02179 | AM | MSK | I020628_160035_340_B.FIT | 4 |  | I | [5 empty] | 1/2 aperature vertical, left masked.
02179 | AM | MSK | I020628_160038_223_B.FIT | 4 |  | I
02179 | AM | MSK | I020628_160042_223_B.FIT | 4 |  | I
02179 | AN | MSK | Focus
02179 | AN | MSK | I020628_163025_473_B.FIT | 4 | 671 | I | 15 | 0.225 | 0 |  |  | Recollect images again, all with the SAME MYLAR MASK (except the last, completely masked image).
02179 | AN | MSK | I020628_163028_323_B.FIT | 4 |  | I | [5 empty] | Full aperature, no mask.
02179 | AN | MSK | I020628_163031_206_B.FIT | 4 |  | I | [5 empty] | Using dark image I020628_135404_290_B.FIT. Source at 12Amp.
02179 | AN | MSK | I020628_163035_056_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163126_973_B.FIT | 4 | 672 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature horizontal, bottom masked (as you look towards source).
02179 | AN | MSK | I020628_163129_806_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163132_706_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163136_706_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163258_523_B.FIT | 4 | 673 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature horizontal, top masked.
02179 | AN | MSK | I020628_163301_356_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163304_256_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163308_256_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163455_356_B.FIT | 4 | 674 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature vertical, right masked.
02179 | AN | MSK | I020628_163458_190_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163501_073_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163505_123_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163616_906_B.FIT | 4 | 675 | I | 15 | 0.225 | 0 |  |  | 1/2 aperature vertical, left masked.
02179 | AN | MSK | I020628_163619_740_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163622_640_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163626_640_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163706_473_B.FIT | 4 | 676 | I | 15 | 0.225 | 0 |  |  | Full aperature, no mask.
02179 | AN | MSK | I020628_163709_323_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163712_206_B.FIT | 4 |  | I
02179 | AN | MSK | I020628_163716_306_B.FIT | 4 |  | I
02179 | AO | DRKW | Darks
02179 | AO | DRKW | I020628_164058_106_B.FIT | 4 | 677 | I | 15 | 0.225 | 0 |  |  | Masked entire aperature. Used wooden box at same location as the mylar.
02179 | AO | DRKW | I020628_164100_940_B.FIT | 4 |  | I | [5 empty] | Note: IRFPA temp is OK.
02179 | AP | MSK | Focus
02179 | AP | MSK | I020628_164103_823_B.FIT | 4 |  | I | [5 empty] | Removed box mask.
02179 | AP | MSK | I020628_164107_690_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171346_423_B.FIT | 4 | 678 | I | 10 | 0.225 | 0 |  |  | Having problems with above boresight test: box photometry varies too much between images.
02179 | AP | MSK | I020628_171349_306_B.FIT | 4 |  | I | [5 empty] | Step towards astigmatic line focus to verify slit is focused on detector.
02179 | AP | MSK | I020628_171352_140_B.FIT | 4 |  | I | [5 empty] | See NO pinholes! Didn't expect to lose pinholes with -5mm movement.
02179 | AP | MSK | I020628_171356_290_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171654_706_B.FIT | 4 | 679 | I | 15 | 0.225 | 0 |  |  | Problem solved: An object was in the way!
02179 | AP | MSK | I020628_171657_556_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171700_440_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171704_540_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171756_273_B.FIT | 4 | 680 | I | 15 | 0.225 | 0 |  |  | Try again at same focus. Got pinholes.
02179 | AP | MSK | I020628_171759_106_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171801_990_B.FIT | 4 |  | I
02179 | AP | MSK | I020628_171806_040_B.FIT | 4 |  | I
02179 | AQ | MSK | Focus
02179 | AQ | MSK | I020628_171927_473_B.FIT | 4 | 681 | I | 10 | 0.225 | 0 |  |  | Proceed to Knife-edge test. Changed focus.
02179 | AQ | MSK | I020628_171930_306_B.FIT | 4 |  | I | [5 empty] | One-pixel "photometry" of X-Slit location vs DN resides at /hammergse/logs_excel/HRI_SIM_DN_vs_XSlit.xls.
02179 | AQ | MSK | I020628_171933_190_B.FIT | 4 |  | I | [5 empty] | This spreadsheet also includes photometry from the "best" focus at 15.0mm (Exp IDs 659 - 664).
02179 | AQ | MSK | I020628_171937_140_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172306_090_B.FIT | 4 | 682 | I | 10 | 0.175 | 0 |  |  | Moved X-slit by 50 microns. Signal barely increased (up 3 of 275 DN).
02179 | AQ | MSK | I020628_172308_940_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172311_823_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172315_823_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172421_306_B.FIT | 4 | 683 | I | 10 | 0.125 | 0 |  |  | Moved X-slit by 50 microns. Signal decreased. Try other direction.
02179 | AQ | MSK | I020628_172424_156_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172427_040_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172431_140_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172654_106_B.FIT | 4 | 684 | I | 10 | 0.275 | 0 |  |  | Moved X-slit.
02179 | AQ | MSK | I020628_172656_940_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172659_840_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172703_840_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172745_556_B.FIT | 4 | 685 | I | 10 | 0.325 | 0 |  |  | Moved X-slit. Signal increased.
02179 | AQ | MSK | I020628_172748_406_B.FIT | 4 |  | I | [5 empty] | Temp (FPA) = ~88K, Temp (Optical Bench) = ~123K
02179 | AQ | MSK | I020628_172751_290_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172755_440_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172846_723_B.FIT | 4 | 686 | I | 10 | 0.375 | 0 |  |  | Moved X-slit. Signal increased.
02179 | AQ | MSK | I020628_172849_556_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172852_440_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172856_306_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172940_323_B.FIT | 4 | 687 | I | 10 | 0.425 | 0 |  |  | Moved X-slit. Signal increased.
02179 | AQ | MSK | I020628_172943_156_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172946_040_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_172950_040_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173135_823_B.FIT | 4 | 688 | I | 10 | 0.475 | 0 |  |  | Moved X-slit.
02179 | AQ | MSK | I020628_173138_656_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173141_540_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173145_490_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173248_123_B.FIT | 4 | 689 | I | 10 | 0.4 | 0 |  |  | Moved X-slit.
02179 | AQ | MSK | I020628_173250_956_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173253_840_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173257_840_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173333_690_B.FIT | 4 | 690 | I | 10 | 0.35 | 0 |  |  | Peak around .380mm.
02179 | AQ | MSK | I020628_173336_540_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173339_423_B.FIT | 4 |  | I
02179 | AQ | MSK | I020628_173343_523_B.FIT | 4 |  | I
02179 | AR | MSK | Focus
02179 | AR | MSK | I020628_174113_590_B.FIT | 4 | 691 | I | 6 | 0.125 | 0 |  |  | Another Knife-edge test. Changed focus and moved X-slit.
02179 | AR | MSK | I020628_174116_440_B.FIT | 4 |  | I | [5 empty] | One-pixel "photometry" of X-Slit location vs DN resides at /hammergse/logs_excel/HRI_SIM_DN_vs_XSlit.xls.
02179 | AR | MSK | I020628_174119_323_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174123_223_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174214_306_B.FIT | 4 | 692 | I | 6 | 0.175 | 0 |  |  | Moved X-Slit. Dim.
02179 | AR | MSK | I020628_174217_156_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174220_040_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174223_940_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174258_256_B.FIT | 4 | 693 | I | 6 | 0.225 | 0 |  |  | Moved X-Slit. Slow increase.
02179 | AR | MSK | I020628_174301_106_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174303_990_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174307_890_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174340_490_B.FIT | 4 | 694 | I | 6 | 0.275 | 0 |  |  | Moved X-Slit. Slow increase.
02179 | AR | MSK | I020628_174343_340_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174346_223_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174350_323_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174439_156_B.FIT | 4 | 695 | I | 6 | 0.325 | 0 |  |  | Moved X-Slit. Slow increase.
02179 | AR | MSK | I020628_174442_006_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174444_890_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174448_790_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174520_756_B.FIT | 4 | 696 | I | 6 | 0.375 | 0 |  |  | Moved X-Slit. Slow increase.
02179 | AR | MSK | I020628_174523_606_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174526_490_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174530_390_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174621_356_B.FIT | 4 | 697 | I | 6 | 0.425 | 0 |  |  | Moved X-Slit. Dramatic decrease.
02179 | AR | MSK | I020628_174624_206_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174627_090_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174630_990_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174707_806_B.FIT | 4 | 698 | I | 6 | 0.475 | 0 |  |  | Moved X-slit.
02179 | AR | MSK | I020628_174710_656_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174713_540_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174717_490_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174756_706_B.FIT | 4 | 699 | I | 6 | 0.35 | 0 |  |  | Moved X-slit.
02179 | AR | MSK | I020628_174759_556_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174802_440_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174806_440_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174838_056_B.FIT | 4 | 700 | I | 6 | 0.4 | 0 |  |  | Moved X-slit.
02179 | AR | MSK | I020628_174840_906_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174843_790_B.FIT | 4 |  | I
02179 | AR | MSK | I020628_174847_690_B.FIT | 4 |  | I
02179 | AS | MSK | Focus
02179 | AS | MSK | I020628_180554_540_B.FIT | 4 | 701 | I | 8 | 0.225 | 0 |  |  | Another Knife-edge test. Changed focus and moved X-slit.
02179 | AS | MSK | I020628_180557_373_B.FIT | 4 |  | I | [5 empty] | One-pixel "photometry" of X-Slit location vs DN resides at /hammergse/logs_excel/HRI_SIM_DN_vs_XSlit.xls.
02179 | AS | MSK | I020628_180600_256_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180604_256_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180651_123_B.FIT | 4 | 702 | I | 8 | 0.275 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_180653_973_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180656_856_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180700_673_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180739_106_B.FIT | 4 | 703 | I | 8 | 0.325 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_180741_940_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180744_823_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180748_773_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180828_723_B.FIT | 4 | 704 | I | 8 | 0.35 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_180831_573_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180834_456_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180838_556_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180912_756_B.FIT | 4 | 705 | I | 8 | 0.375 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_180915_590_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180918_490_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_180922_240_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181000_406_B.FIT | 4 | 706 | I | 8 | 0.4 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_181003_240_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181006_140_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181010_140_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181046_623_B.FIT | 4 | 707 | I | 8 | 0.425 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_181049_473_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181052_356_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181056_456_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181135_640_B.FIT | 4 | 708 | I | 8 | 0.475 | 0 |  |  | Moved X-slit.
02179 | AS | MSK | I020628_181138_473_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181141_356_B.FIT | 4 |  | I
02179 | AS | MSK | I020628_181145_406_B.FIT | 4 |  | I
02179 | AT | MSK | Focus
02179 | AT | MSK | I020628_181546_006_B.FIT | 4 | 709 | I | 12.5 | 0.225 | 0 |  |  | Another Knife-edge test. Changed focus and moved X-slit.
02179 | AT | MSK | I020628_181548_840_B.FIT | 4 |  | I | [5 empty] | One-pixel "photometry" of X-Slit location vs DN resides at /hammergse/logs_excel/HRI_SIM_DN_vs_XSlit.xls.
02179 | AT | MSK | I020628_181551_740_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181555_740_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181634_240_B.FIT | 4 | 710 | I | 12.5 | 0.275 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_181637_073_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181639_956_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181644_006_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181722_790_B.FIT | 4 | 711 | I | 12.5 | 0.325 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_181725_623_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181728_506_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181732_506_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181810_340_B.FIT | 4 | 712 | I | 12.5 | 0.35 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_181813_173_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181816_056_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181820_156_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181858_540_B.FIT | 4 | 713 | I | 12.5 | 0.375 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_181901_373_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181904_256_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181908_123_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181945_856_B.FIT | 4 | 714 | I | 12.5 | 0.4 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_181948_690_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181951_573_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_181955_590_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182033_306_B.FIT | 4 | 715 | I | 12.5 | 0.425 | 0 |  |  | Moved X-slit.
02179 | AT | MSK | I020628_182036_140_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182039_023_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182043_040_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182124_006_B.FIT | 4 | 716 | I | 12.5 | 0.475 | 0 |  |  | Moved X-slit. Shut down at 10mm
02179 | AT | MSK | I020628_182126_840_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182129_740_B.FIT | 4 |  | I
02179 | AT | MSK | I020628_182133_740_B.FIT | 4 |  | I

## 020629
YRDAY | SEQ | TYPE | Filenames for 02180 (June 29 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Van Cleve
|  |  | Optical Bench Temp: | not recorded
564449 - Collected IRFPA Image Table
02180 | AA | LVL | Level
02180 | AA | LVL | I020629_110217_923_B.FIT | 6 | 701 | I | [5 empty] | Dark. Reference = wood box in corner.
02180 | AA | LVL | I020629_110221_490_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | Allmode dark
02180 | AB | DRKW | I020629_110604_023_B.FIT | 1 | 702 | I | [5 empty] | Dark. Temp (FPA) = ~xxxK, Temp (Optical Bench) = ~xxxK
02180 | AB | DRKW | I020629_110606_806_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110609_773_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110612_540_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110615_490_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110618_256_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110621_223_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110625_073_B.FIT | 1 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110752_106_B.FIT | 2 | 703 | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110753_590_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110755_073_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110756_373_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110757_840_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110759_323_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110800_806_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110803_406_B.FIT | 2 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110928_323_B.FIT | 3 | 704 | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110929_056_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110929_790_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110930_540_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110931_273_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110932_023_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110932_723_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_110934_740_B.FIT | 3 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111104_173_B.FIT | 4 | 705 | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111107_006_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111109_906_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111113_706_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111224_706_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111227_556_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111230_440_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111234_540_B.FIT | 4 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111403_706_B.FIT | 5 | 706 | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111406_523_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111409_440_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111412_256_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111415_173_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111417_990_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111420_756_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111423_723_B.FIT | 5 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111551_906_B.FIT | 6 | 707 | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111555_523_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111559_006_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111602_506_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111606_090_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111609_640_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111613_190_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AB | DRKW | I020629_111616_590_B.FIT | 6 |  | I | [5 empty] | Dark
02180 | AC | CHK | Memory check
02180 | AC | CHK | I020629_112247_506_B.FIT | 7 | 708 | I | [5 empty] | Memory check mode - No data.
02180 | AC | CHK | I020629_112250_373_B.FIT | 7 |  | I | [5 empty] | Memory check mode - No data.
02180 | AD | STIMIR | Stimulator
02180 | AD | STIMIR | I020629_115042_306_B.FIT | 4 | 709 | I | [5 empty] | Waiting for Dennis. Decided to rehearse calibration. Stimulator turned on. DAC at X60. Box in corner.
02180 | AD | STIMIR | I020629_115045_156_B.FIT | 4 |  | I
02180 | AD | STIMIR | I020629_115048_040_B.FIT | 4 |  | I
02180 | AD | STIMIR | I020629_115051_940_B.FIT | 4 |  | I
02180 | AD | STIMIR | I020629_115836_723_B.FIT | 4 | 710 | I | [5 empty] | Stimulator still on. DAC at X70. Box in corner.
02180 | AD | STIMIR | I020629_115839_606_B.FIT | 4 |  | I
02180 | AD | STIMIR | I020629_115842_440_B.FIT | 4 |  | I
02180 | AD | STIMIR | I020629_115846_690_B.FIT | 4 |  | I
02180 | AE | MSK | Focus
02180 | AE | MSK | I020629_120435_223_B.FIT | 4 | 711 | I | 5 | 0 | 0 |  |  | Dennis is back. Stimulator off. See if we can see source at yesterday's position, 15mm focus, 380 Kl
02180 | AE | MSK | I020629_120438_073_B.FIT | 4 |  | I | [5 empty] | Focus is OK. However, got slightly diagonal ghosts (+ slope, from left to right in image) above and below each pinhole spectrum.
02180 | AE | MSK | I020629_120440_956_B.FIT | 4 |  | I | [5 empty] | Move source in X-slit direction to see if ghosts are artifacts of the source.
02180 | AE | MSK | I020629_120445_006_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_121144_523_B.FIT | 4 | 712 | I | 5 | 0.05 | 0 |  |  | Move X-Slit to +50 micron
02180 | AE | MSK | I020629_121147_356_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_121150_256_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_121154_256_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122015_306_B.FIT | 4 | 713 | I | 5 | -0.05 | 0 |  |  | Move X-Slit to -50 micron
02180 | AE | MSK | I020629_122018_140_B.FIT | 4 |  | I | [5 empty] | The + slope that increases ~linearly with increasing wavelength. SIM should be in focus.
02180 | AE | MSK | I020629_122021_023_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122025_090_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122223_123_B.FIT | 4 | 714 | I | 5 | -0.1 | 0 |  |  | Move X-Slit to -100 micron. No improvement. Try 100 micron steps.
02180 | AE | MSK | I020629_122225_956_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122228_856_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122232_856_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122446_523_B.FIT | 4 | 715 | I | 5 | -0.2 | 0 |  |  | Move X-Slit to -200 micron. No big improvement.
02180 | AE | MSK | I020629_122449_373_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122452_256_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122456_156_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122850_473_B.FIT | 4 | 716 | I | 5 | 0.1 | 0 |  |  | Move X-Slit to +100 micron.
02180 | AE | MSK | I020629_122853_306_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122856_206_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_122900_206_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_123002_923_B.FIT | 4 | 717 | I | 5 | 0.2 | 0 |  |  | Moved X-Slit to +200 micron. No big improvement. The -50 micron position seems to be the best of this set.
02180 | AE | MSK | I020629_123005_773_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_123008_656_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_123012_556_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_123657_356_B.FIT | 4 | 718 | I | 5 | -0.05 | 0 |  |  | Adjusted source and moved X-slit to -50 micron.
02180 | AE | MSK | I020629_123700_190_B.FIT | 4 |  | I | [5 empty] | Image improved: Ghost wings are ~gone. Image shifted up by 10 pixels.
02180 | AE | MSK | I020629_123703_073_B.FIT | 4 |  | I
02180 | AE | MSK | I020629_123706_890_B.FIT | 4 |  | I
02180 | AF | MSK | Focus
02180 | AF | MSK | I020629_124309_823_B.FIT | 4 | 719 | I | 0 | 0 | 0 |  |  | Set everything to 0, including the focus.
02180 | AF | MSK | I020629_124312_656_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_124315_540_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_124319_590_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125113_806_B.FIT | 4 | 720 | I | 0 | 0 | 0 |  |  | Realigned concentrating optics.
02180 | AF | MSK | I020629_125116_640_B.FIT | 4 |  | I | [5 empty] | DN increased such that only off by a factor of ~3 from yesterday. Image shifted ~10 pixels down.
02180 | AF | MSK | I020629_125119_523_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125123_573_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125423_223_B.FIT | 4 | 721 | I | 0 | 0.05 | 0 |  |  | Let's sweep at this position.
02180 | AF | MSK | I020629_125426_106_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125428_940_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125433_090_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125528_356_B.FIT | 4 | 722 | I | 0 | -0.05 | 0 |  |  | +1600. Best?
02180 | AF | MSK | I020629_125531_206_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125534_090_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_125538_240_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_130115_556_B.FIT | 4 | 723 | I | 5 | -0.05 | 0 |  |  | Signal is down and wings are back.
02180 | AF | MSK | I020629_130118_406_B.FIT | 4 |  | I | [5 empty] | FWHM = 3.4 at column 888.
02180 | AF | MSK | I020629_130121_290_B.FIT | 4 |  | I | [5 empty] | Let's scan horizontally.
02180 | AF | MSK | I020629_130125_290_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131205_223_B.FIT | 4 | 724 | I | 5 | -0.15 | 0 |  |  | Moved X-Slit to +50 micron.
02180 | AF | MSK | I020629_131208_056_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131210_956_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131214_706_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131303_406_B.FIT | 4 | 725 | I | 5 | 0.05 | 0 |  |  | Moved X-Slit to +50 micron.
02180 | AF | MSK | I020629_131306_256_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131309_140_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131313_190_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131509_590_B.FIT | 4 | 726 | I | 5 | 0.15 | 0 |  |  | Moved X-Slit to +150 micron.
02180 | AF | MSK | I020629_131512_423_B.FIT | 4 |  | I | [5 empty] | Improved.
02180 | AF | MSK | I020629_131515_323_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131519_323_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131625_706_B.FIT | 4 | 727 | I | 5 | 0.25 | 0 |  |  | Moved X-Slit to +250 micron.
02180 | AF | MSK | I020629_131628_556_B.FIT | 4 |  | I | [5 empty] | Improved. Getting closer to yesterday's brightness.
02180 | AF | MSK | I020629_131631_440_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131635_540_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131758_006_B.FIT | 4 | 728 | I | 5 | 0.35 | 0 |  |  | Moved X-Slit to +350 micron.
02180 | AF | MSK | I020629_131800_856_B.FIT | 4 |  | I | [5 empty] | Brightness went down.
02180 | AF | MSK | I020629_131803_740_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131807_790_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_131913_506_B.FIT | 4 | 729 | I | 5 | 0.3 | 0 |  |  | Moved X-Slit to +300 micron.
02180 | AF | MSK | I020629_131916_356_B.FIT | 4 |  | I | [5 empty] | Brightness went down. +250 is close to best for X-Slit.
02180 | AF | MSK | I020629_131919_240_B.FIT | 4 |  | I | [5 empty] | Source = 12 Amp.
02180 | AF | MSK | I020629_131923_090_B.FIT | 4 |  | I
02180 | AF | MSK | I020629_132658_856_B.FIT | 4 | 730 | I | 5 | 0.25 | 0 |  |  | Moved X-Slit to +250 micron, best intensity for X-Slit.
02180 | AF | MSK | I020629_132701_690_B.FIT | 4 |  | I | [5 empty] | Source = 15 Amp, for best performance at short wavelengths. Trying to best spatial focus.
02180 | AF | MSK | I020629_132704_590_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 4.72.
02180 | AF | MSK | I020629_132708_340_B.FIT | 4 |  | I
02180 | AG | MSK | Focus
02180 | AG | MSK | I020629_133336_940_B.FIT | 4 | 731 | I | 2 | 0 | 0 |  |  | Rezero Z-focus and X-Slit for previous values (5.000mm and 0.250mm, respectively).
02180 | AG | MSK | I020629_133339_773_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 4.96, measured at the center pinhole
02180 | AG | MSK | I020629_133342_673_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.8, measured at the center pinhole.
02180 | AG | MSK | I020629_133346_673_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_134021_440_B.FIT | 4 | 732 | I | 4 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_134024_273_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 3.64; an improvement.
02180 | AG | MSK | I020629_134027_173_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.21
02180 | AG | MSK | I020629_134031_173_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_134330_256_B.FIT | 4 | 733 | I | 6 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_134333_140_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 4.17; Widened.
02180 | AG | MSK | I020629_134335_973_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.62
02180 | AG | MSK | I020629_134340_123_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_134532_990_B.FIT | 4 | 734 | I | 8 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_134535_823_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 3.20
02180 | AG | MSK | I020629_134538_723_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 2.96
02180 | AG | MSK | I020629_134542_473_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_134733_740_B.FIT | 4 | 735 | I | -2 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_134736_590_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 5.72
02180 | AG | MSK | I020629_134739_473_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.36
02180 | AG | MSK | I020629_134743_223_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_134934_106_B.FIT | 4 | 736 | I | -4 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_134936_940_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 6.02
02180 | AG | MSK | I020629_134939_840_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.66
02180 | AG | MSK | I020629_134943_840_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_135035_740_B.FIT | 4 | 737 | I | -6 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_135038_573_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 6.41
02180 | AG | MSK | I020629_135041_456_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.79
02180 | AG | MSK | I020629_135045_473_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_135129_990_B.FIT | 4 | 738 | I | -8 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_135132_823_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 7.67
02180 | AG | MSK | I020629_135135_723_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.47
02180 | AG | MSK | I020629_135139_723_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_140213_106_B.FIT | 4 | 739 | I | 6 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_140215_940_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 4.01
02180 | AG | MSK | I020629_140218_823_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.55
02180 | AG | MSK | I020629_140222_923_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_140329_206_B.FIT | 4 | 740 | I | 10 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_140332_040_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 2.88
02180 | AG | MSK | I020629_140334_923_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 2.72
02180 | AG | MSK | I020629_140338_973_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_141207_223_B.FIT | 4 | 741 | I | 12 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_141210_073_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 4.3
02180 | AG | MSK | I020629_141212_956_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 3.36
02180 | AG | MSK | I020629_141217_056_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_141659_706_B.FIT | 4 | 742 | I | 14 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_141702_540_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 7.1
02180 | AG | MSK | I020629_141705_423_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 2.6
02180 | AG | MSK | I020629_141709_373_B.FIT | 4 |  | I
02180 | AG | MSK | I020629_142520_323_B.FIT | 4 | 743 | I | 20 | 0 | 0 |  |  | Moved Z-focus.
02180 | AG | MSK | I020629_142523_206_B.FIT | 4 |  | I | [5 empty] | At 1.33 micron (col 270), FWHM = 5.8
02180 | AG | MSK | I020629_142526_040_B.FIT | 4 |  | I | [5 empty] | At 3.90 micron (col 888), FWHM = 2.9 Lunch….
02180 | AG | MSK | I020629_142530_106_B.FIT | 4 |  | I
02180 | AH | DRKW | Darks
02180 | AH | DRKW | I020629_154754_623_B.FIT | 4 | 750 | I | n/a | n/a | n/a |  |  | Darks.
02180 | AH | DRKW | I020629_154759_056_B.FIT | 4 |  | I
02180 | AH | DRKW | I020629_154933_806_B.FIT | 4 | 751 | I | n/a | n/a | n/a |  |  | Darks. Use this dark (I020629_154933_806_B.FIT) for preceeding analysis.
02180 | AH | DRKW | I020629_154936_656_B.FIT | 4 |  | I
02180 | AH | DRKW | I020629_154939_540_B.FIT | 4 |  | I
02180 | AH | DRKW | I020629_154943_740_B.FIT | 4 |  | I
02180 | AI | MSK | Focus
02180 | AI | MSK | I020629_160318_873_B.FIT | 4 | 752 | I | 20 | 0 | 0 |  |  | Redo pre-lunch image using best focus (10mm). Source = 15 Amp.
02180 | AI | MSK | I020629_160321_756_B.FIT | 4 |  | I | [5 empty] | Oops, lost the image. Use 100 micron steps to find the image.
02180 | AI | MSK | I020629_160324_590_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160328_656_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160737_340_B.FIT | 4 | 753 | I | 20 | 0.1 | 0 |  |  | Image hunt #1...
02180 | AI | MSK | I020629_160740_173_B.FIT | 4 |  | I | [5 empty] | See some image.
02180 | AI | MSK | I020629_160743_056_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160747_106_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160853_740_B.FIT | 4 | 754 | I | 20 | 0.2 | 0 |  |  | Image hunt #2...
02180 | AI | MSK | I020629_160856_573_B.FIT | 4 |  | I | [5 empty] | Getting closer. Right direction but signal is low.
02180 | AI | MSK | I020629_160859_456_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160903_506_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_160959_223_B.FIT | 4 | 755 | I | 20 | 0.3 | 0 |  |  | Image hunt #3...
02180 | AI | MSK | I020629_161002_073_B.FIT | 4 |  | I | [5 empty] | Getting closer. Signal brighter.
02180 | AI | MSK | I020629_161004_956_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161008_956_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161110_523_B.FIT | 4 | 756 | I | 20 | 0.4 | 0 |  |  | Image hunt #4...
02180 | AI | MSK | I020629_161113_373_B.FIT | 4 |  | I | [5 empty] | Image looks blurry and signal is lower.
02180 | AI | MSK | I020629_161116_256_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161120_156_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161325_156_B.FIT | 4 | 757 | I | 20 | -0.2 | 0 |  |  | Image hunt #5…
02180 | AI | MSK | I020629_161328_006_B.FIT | 4 |  | I | [5 empty] | See nothing.
02180 | AI | MSK | I020629_161330_890_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161334_790_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161428_406_B.FIT | 4 | 758 | I | 20 | 0.5 | 0 |  |  | Image hunt #6…
02180 | AI | MSK | I020629_161431_240_B.FIT | 4 |  | I | [5 empty] | Image blurry and signal decreasing at short wavelengths (ie, 2.2 microns)
02180 | AI | MSK | I020629_161434_140_B.FIT | 4 |  | I | [5 empty] | Maybe on a diffraction ring? Keep increasing X-Slit.
02180 | AI | MSK | I020629_161437_940_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161725_573_B.FIT | 4 | 759 | I | 20 | 0.6 | 0 |  |  | Signal is increasing. Must have been on a diffraction ring.
02180 | AI | MSK | I020629_161728_423_B.FIT | 4 |  | I | [5 empty] | DN 61
02180 | AI | MSK | I020629_161731_306_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161735_456_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161854_073_B.FIT | 4 | 760 | I | 20 | 0.7 | 0 |  |  | Signal is about the same as last image.
02180 | AI | MSK | I020629_161856_923_B.FIT | 4 |  | I | [5 empty] | DN 58
02180 | AI | MSK | I020629_161859_806_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_161904_006_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_162004_623_B.FIT | 4 | 761 | I | 20 | 0.8 | 0 |  |  | Image is brighter and sharper.
02180 | AI | MSK | I020629_162007_473_B.FIT | 4 |  | I | [5 empty] | DN 82
02180 | AI | MSK | I020629_162010_356_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_162014_256_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_162115_523_B.FIT | 4 | 762 | I | 20 | 0.9 | 0 |  |  | Image still brighter
02180 | AI | MSK | I020629_162118_373_B.FIT | 4 |  | I | [5 empty] | DN 114
02180 | AI | MSK | I020629_162121_256_B.FIT | 4 |  | I
02180 | AI | MSK | I020629_162125_256_B.FIT | 4 |  | I
02180 | AJ | MSK | Focus
02180 | AJ | MSK | I020629_162249_073_B.FIT | 4 | 763 | I | 20 | 1 | 0
02180 | AJ | MSK | I020629_162251_923_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162254_806_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162258_706_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162734_756_B.FIT | 4 | 764 | I | 20 | 1.1 | 0
02180 | AJ | MSK | I020629_162737_606_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162740_490_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162744_390_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162817_006_B.FIT | 4 | 765 | I | 20 | 1.2 | 0
02180 | AJ | MSK | I020629_162819_890_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162822_723_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162826_873_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162914_256_B.FIT | 4 | 766 | I | 20 | 1.3 | 0
02180 | AJ | MSK | I020629_162917_090_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162919_990_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_162923_940_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_163023_873_B.FIT | 4 | 767 | I | 20 | 1.4 | 0
02180 | AJ | MSK | I020629_163026_723_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_163029_606_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_163033_506_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164004_323_B.FIT | 4 | 768 | I | 20 | 1.5 | 0 |  |  | DN 2796 (big jump)
02180 | AJ | MSK | I020629_164007_156_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164010_040_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164013_906_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164044_506_B.FIT | 4 | 769 | I | 20 | 1.6 | 0
02180 | AJ | MSK | I020629_164047_340_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164050_223_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164054_223_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164134_923_B.FIT | 4 | 770 | I | 20 | 1.7 | 0
02180 | AJ | MSK | I020629_164137_756_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164140_656_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164144_406_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164223_856_B.FIT | 4 | 771 | I | 20 | 1.8 | 0
02180 | AJ | MSK | I020629_164226_690_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164229_573_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164233_390_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164548_890_B.FIT | 4 | 772 | I | 20 | 1.9 | 0
02180 | AJ | MSK | I020629_164551_740_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164554_623_B.FIT | 4 |  | I
02180 | AJ | MSK | I020629_164558_723_B.FIT | 4 |  | I
02180 | AK | MSK | Focus
02180 | AK | MSK | I020629_164653_523_B.FIT | 4 | 773 | I | 20 | 2 | 0
02180 | AK | MSK | I020629_164656_373_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_164659_256_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_164703_256_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165133_423_B.FIT | 4 | 774 | I | 20 | 2.1 | 0
02180 | AK | MSK | I020629_165136_273_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165139_156_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165143_156_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165212_773_B.FIT | 4 | 775 | I | 20 | 2.2 | 0 |  |  | BOGUS set!
02180 | AK | MSK | I020629_165215_623_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165218_506_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165222_206_B.FIT | 4 |  | I
02180 | AK | MSK | i020629_165258_123_b.FIT | 4 | 776 | I | 20 | 2.2 | 0 |  |  | Good set
02180 | AK | MSK | I020629_165300_973_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165304_573_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165307_940_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165401_023_B.FIT | 4 | 777 | I | 20 | 2.3 | 0
02180 | AK | MSK | I020629_165403_873_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165406_756_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_165410_756_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_171003_856_B.FIT | 4 | 778 | I | 15 | 1.8 | 0
02180 | AK | MSK | I020629_171006_706_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_171009_590_B.FIT | 4 |  | I
02180 | AK | MSK | I020629_171013_690_B.FIT | 4 |  | I
02180 | AL | MSK | Focus
02180 | AL | MSK | I020629_171511_823_B.FIT | 4 | 779 | I | 12 | 1.8 | 0
02180 | AL | MSK | I020629_171516_606_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_171521_356_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_171526_573_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_171757_373_B.FIT | 4 | 780 | I | 10 | 1.8 | 0
02180 | AL | MSK | I020629_171800_806_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_171807_490_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_171813_156_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172013_073_B.FIT | 4 | 781 | I | 10 | 1.7 | 0 |  |  | Start knife edge.
02180 | AL | MSK | I020629_172018_490_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172025_056_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172030_423_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172115_290_B.FIT | 4 | 782 | I | 10 | 1.6 | 0
02180 | AL | MSK | I020629_172121_873_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172128_673_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172135_590_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172206_190_B.FIT | 4 | 783 | I | 10 | 1.5 | 0
02180 | AL | MSK | I020629_172211_840_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172219_040_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172222_473_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172301_823_B.FIT | 4 | 784 | I | 10 | 1.4 | 0
02180 | AL | MSK | I020629_172305_256_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172309_406_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172312_923_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172350_323_B.FIT | 4 | 785 | I | 10 | 1.3 | 0
02180 | AL | MSK | I020629_172354_206_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172358_690_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172402_406_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172440_256_B.FIT | 4 | 786 | I | 10 | 1.2 | 0 |  |  | Signal dropped dramatically.
02180 | AL | MSK | I020629_172444_523_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172447_956_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172451_073_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172702_156_B.FIT | 4 | 787 | I | 10 | 1.9 | 0 |  |  | Signal continues to drop.
02180 | AL | MSK | I020629_172705_006_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172707_890_B.FIT | 4 |  | I
02180 | AL | MSK | I020629_172711_990_B.FIT | 4 |  | I
02180 | AM | MSK | Focus
02180 | AM | MSK | I020629_172759_740_B.FIT | 4 | 788 | I | 10 | 2 | 0
02180 | AM | MSK | I020629_172802_590_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172805_473_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172809_473_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172849_673_B.FIT | 4 | 789 | I | 10 | 2.1 | 0
02180 | AM | MSK | I020629_172853_623_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172857_390_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172900_723_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172944_290_B.FIT | 4 | 790 | I | 10 | 2.2 | 0
02180 | AM | MSK | I020629_172947_140_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172950_023_B.FIT | 4 |  | I
02180 | AM | MSK | I020629_172954_023_B.FIT | 4 |  | I
02180 | AN | MSK | Focus
02180 | AN | MSK | I020629_173816_256_B.FIT | 4 | 791 | I | 5 | 1.8 | 0 |  |  | Changed Z-Focus.
02180 | AN | MSK | I020629_173819_140_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_173821_973_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_173826_040_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_173908_290_B.FIT | 4 | 792 | I | 5 | 1.7 | 0
02180 | AN | MSK | I020629_173911_123_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_173914_023_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_173918_023_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174001_106_B.FIT | 4 | 793 | I | 5 | 1.6 | 0
02180 | AN | MSK | I020629_174003_956_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174006_840_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174010_940_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174050_240_B.FIT | 4 | 794 | I | 5 | 1.5 | 0
02180 | AN | MSK | I020629_174053_073_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174055_973_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174059_923_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174258_756_B.FIT | 4 | 795 | I | 5 | 1.4 | 0
02180 | AN | MSK | I020629_174301_606_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174304_490_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174308_390_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174354_756_B.FIT | 4 | 796 | I | 5 | 1.9 | 0
02180 | AN | MSK | I020629_174357_590_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174400_490_B.FIT | 4 |  | I
02180 | AN | MSK | I020629_174404_490_B.FIT | 4 |  | I
02180 | AO | MSK | Focus
02180 | AO | MSK | I020629_174444_156_B.FIT | 4 | 797 | I | 5 | 2 | 0
02180 | AO | MSK | I020629_174447_006_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174449_890_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174453_790_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174539_606_B.FIT | 4 | 798 | I | 5 | 2.1 | 0
02180 | AO | MSK | I020629_174542_456_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174545_340_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174549_540_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174634_290_B.FIT | 4 | 799 | I | 5 | 2.2 | 0
02180 | AO | MSK | I020629_174638_240_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174641_440_B.FIT | 4 |  | I
02180 | AO | MSK | I020629_174644_456_B.FIT | 4 |  | I
02180 | AP | MSK | Focus
02180 | AP | MSK | I020629_174912_456_B.FIT | 4 | 800 | I | 0 | 1.8 | 0 |  |  | Changed Z-Focus.
02180 | AP | MSK | I020629_174915_306_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_174918_190_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_174922_190_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175042_040_B.FIT | 4 | 801 | I | 0 | 1.75 | 0
02180 | AP | MSK | I020629_175044_890_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175047_773_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175051_673_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175136_806_B.FIT | 4 | 802 | I | 0 | 1.7 | 0
02180 | AP | MSK | I020629_175140_023_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175143_806_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175146_873_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175232_590_B.FIT | 4 | 803 | I | 0 | 1.65 | 0
02180 | AP | MSK | I020629_175235_440_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175238_323_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175242_223_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175323_056_B.FIT | 4 | 804 | I | 0 | 1.6 | 0
02180 | AP | MSK | I020629_175325_906_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175328_790_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175332_690_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175416_390_B.FIT | 4 | 805 | I | 0 | 1.85 | 0
02180 | AP | MSK | I020629_175419_240_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175422_123_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175426_023_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175521_290_B.FIT | 4 | 806 | I | 0 | 1.9 | 0
02180 | AP | MSK | I020629_175524_140_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175527_023_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175531_123_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175717_390_B.FIT | 4 | 807 | I | 0 | 1.95 | 0
02180 | AP | MSK | I020629_175720_273_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175723_106_B.FIT | 4 |  | I
02180 | AP | MSK | I020629_175727_306_B.FIT | 4 |  | I
02180 | AQ | MSK | Focus
02180 | AQ | MSK | I020629_175806_556_B.FIT | 4 | 808 | I | 0 | 2 | 0
02180 | AQ | MSK | I020629_175809_406_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_175812_290_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_175816_140_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180210_840_B.FIT | 4 | 809 | I | 0 | 1.55 | 0
02180 | AQ | MSK | I020629_180213_690_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180216_573_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180220_673_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180850_623_B.FIT | 4 | 810 | I | 0 | 1.5 | 0
02180 | AQ | MSK | I020629_180853_473_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180856_356_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180900_256_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180943_640_B.FIT | 4 | 811 | I | 0 | 1.45 | 0
02180 | AQ | MSK | I020629_180946_473_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180949_356_B.FIT | 4 |  | I
02180 | AQ | MSK | I020629_180953_223_B.FIT | 4 |  | I
02180 | AR | MSK | Focus
02180 | AR | MSK | I020629_181132_306_B.FIT | 4 | 812 | I | -5 | 1.7 | 0 |  |  | Changed Z-Focus.
02180 | AR | MSK | I020629_181135_140_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181138_040_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181141_790_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181334_706_B.FIT | 4 | 813 | I | -5 | 1.65 | 0
02180 | AR | MSK | I020629_181337_556_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181340_440_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181344_390_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181510_506_B.FIT | 4 | 814 | I | -5 | 1.6 | 0
02180 | AR | MSK | I020629_181513_340_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181516_240_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181520_240_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181917_056_B.FIT | 4 | 815 | I | -5 | 1.55 | 0
02180 | AR | MSK | I020629_181919_906_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181922_790_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_181926_740_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182018_890_B.FIT | 4 | 816 | I | -5 | 1.5 | 0
02180 | AR | MSK | I020629_182022_223_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182024_606_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182028_606_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182105_956_B.FIT | 4 | 817 | I | -5 | 1.45 | 0
02180 | AR | MSK | I020629_182108_806_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182111_690_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182115_690_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182159_573_B.FIT | 4 | 818 | I | -5 | 1.75 | 0
02180 | AR | MSK | I020629_182202_456_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182205_290_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182209_356_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182250_206_B.FIT | 4 | 819 | I | -5 | 1.8 | 0
02180 | AR | MSK | I020629_182253_040_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182255_940_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182259_940_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182606_890_B.FIT | 4 | 820 | I | -5 | 1.9 | 0
02180 | AR | MSK | I020629_182609_723_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182612_623_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182616_423_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182740_906_B.FIT | 4 | 821 | I | -5 | 1.95 | 0
02180 | AR | MSK | I020629_182743_790_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182746_623_B.FIT | 4 |  | I
02180 | AR | MSK | I020629_182750_823_B.FIT | 4 |  | I
02180 | AS | MSK | Focus
02180 | AS | MSK | I020629_182852_990_B.FIT | 4 | 822 | I | -5 | 2 | 0
02180 | AS | MSK | I020629_182855_823_B.FIT | 4 |  | I
02180 | AS | MSK | I020629_182858_706_B.FIT | 4 |  | I
02180 | AS | MSK | I020629_182902_706_B.FIT | 4 |  | I
02180 | AS | MSK | I020629_182947_590_B.FIT | 4 | 823 | I | -5 | 1.85 | 0
02180 | AS | MSK | I020629_182950_423_B.FIT | 4 |  | I
02180 | AS | MSK | I020629_182953_306_B.FIT | 4 |  | I
02180 | AS | MSK | I020629_182957_306_B.FIT | 4 |  | I
02180 | AT | MSK | Focus
02180 | AT | MSK | I020629_183120_073_B.FIT | 4 | 824 | I | -10 | 1.8 | 0 |  |  | Changed Z-Focus.
02180 | AT | MSK | I020629_183122_923_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183125_806_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183129_706_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183211_356_B.FIT | 4 | 825 | I | -10 | 1.7 | 0
02180 | AT | MSK | I020629_183215_540_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183218_206_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183220_990_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183333_840_B.FIT | 4 | 826 | I | -10 | 1.6 | 0
02180 | AT | MSK | I020629_183336_690_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183339_573_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183343_390_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183443_623_B.FIT | 4 | 827 | I | -10 | 1.5 | 0
02180 | AT | MSK | I020629_183446_473_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183449_356_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183453_256_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183602_490_B.FIT | 4 | 828 | I | -10 | 1.4 | 0
02180 | AT | MSK | I020629_183605_340_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183608_223_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183612_123_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183747_323_B.FIT | 2 | 829 | I | -10 | 1.9 | 0 |  |  | Wrong mode number.
02180 | AT | MSK | I020629_183820_856_B.FIT | 4 | 830 | I | -10 | 1.9 | 0
02180 | AT | MSK | I020629_183823_690_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183826_573_B.FIT | 4 |  | I
02180 | AT | MSK | I020629_183830_390_B.FIT | 4 |  | I
02180 | AU | MSK | Focus
02180 | AU | MSK | I020629_185257_823_B.FIT | 4 | 831 | I | -20 | 1.7 | 0 |  |  | Changed Z-Focus.
02180 | AU | MSK | I020629_185300_673_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185303_556_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185307_756_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185429_073_B.FIT | 4 | 832 | I | -20 | 1.6 | 0
02180 | AU | MSK | I020629_185431_923_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185434_806_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185439_006_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185528_323_B.FIT | 4 | 833 | I | -20 | 1.5 | 0
02180 | AU | MSK | I020629_185531_173_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185534_056_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185538_156_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185719_840_B.FIT | 4 | 834 | I | -20 | 1.4 | 0
02180 | AU | MSK | I020629_185722_690_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185725_573_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185729_423_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185807_290_B.FIT | 4 | 835 | I | -20 | 1.3 | 0
02180 | AU | MSK | I020629_185810_140_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185813_590_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185817_306_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185902_006_B.FIT | 4 | 836 | I | -20 | 1.8 | 0
02180 | AU | MSK | I020629_185904_840_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185907_723_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_185911_673_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190045_873_B.FIT | 4 | 837 | I | -20 | 1.9 | 0
02180 | AU | MSK | I020629_190048_706_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190051_590_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190055_606_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190137_323_B.FIT | 4 | 838 | I | -20 | 2 | 0
02180 | AU | MSK | I020629_190140_156_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190143_040_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190147_056_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190237_840_B.FIT | 4 | 839 | I
02180 | AU | MSK | I020629_190241_523_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190245_306_B.FIT | 4 |  | I
02180 | AU | MSK | I020629_190247_773_B.FIT | 4 |  | I
02180 | AV | STIMIR | Stimulator
02180 | AV | STIMIR | I020629_192139_590_B.FIT | 4 | 900 | I | -20 | 2.1 | 0 |  |  | Turned stimulator on. DAC Hex 80
02180 | AV | STIMIR | I020629_192142_440_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_192145_323_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_192149_373_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_192520_673_B.FIT | 4 | 901 | I | -20 | 2.1 | 0 |  |  | DAC Hex 00
02180 | AV | STIMIR | I020629_192523_523_B.FIT | 4 |  | I | [5 empty] | Subtracting STIM images produces spectra.
02180 | AV | STIMIR | I020629_192526_406_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_192530_356_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193410_790_B.FIT | 4 | 902 | I | -20 | 2.1 | 0 |  |  | DAC Hex 90
02180 | AV | STIMIR | I020629_193413_623_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193416_506_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193420_606_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193905_940_B.FIT | 4 | 903 | I | -20 | 2.1 | 0 |  |  | DAC Hex A0
02180 | AV | STIMIR | I020629_193908_773_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193911_656_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_193915_606_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_194212_740_B.FIT | 4 | 904 | I | -20 | 2.1 | 0 |  |  | DAC Hex 00
02180 | AV | STIMIR | I020629_194215_573_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_194218_473_B.FIT | 4 |  | I
02180 | AV | STIMIR | I020629_194222_273_B.FIT | 4 |  | I
02180 | AW | AL | Argon lamp
02180 | AW | AL | I020629_194606_140_B.FIT | 4 | 905 | I | -20 | 2.1 | 0 |  |  | Source = Argon lamp. Slit = 1mm
02180 | AW | AL | I020629_194608_990_B.FIT | 4 |  | I | [5 empty] | Signal low but GOT LINES!
02180 | AW | AL | I020629_194611_873_B.FIT | 4 |  | I
02180 | AW | AL | I020629_194615_873_B.FIT | 4 |  | I
02180 | AW | AL | I020629_195650_523_B.FIT | 4 | 906 | I | -20 | 2.1 | 0 |  |  | Move Argon lamp and repeat.
02180 | AW | AL | I020629_195654_023_B.FIT | 4 |  | I | [5 empty] | Signal decreased dramatically; Lamp off slit, so try again.
02180 | AW | AL | I020629_195656_256_B.FIT | 4 |  | I
02180 | AW | AL | I020629_195700_356_B.FIT | 4 |  | I
02180 | AW | AL | I020629_195938_823_B.FIT | 4 | 907 | I | -20 | 2.1 | 0 |  |  | Changed Argon lamp location.
02180 | AW | AL | I020629_195941_673_B.FIT | 4 |  | I | [5 empty] | Got lines again.
02180 | AW | AL | I020629_195944_556_B.FIT | 4 |  | I
02180 | AW | AL | I020629_195948_556_B.FIT | 4 |  | I
02180 | AW | AL | I020629_200301_623_B.FIT | 4 | 908 | I | -20 | 2.1 | 0 |  |  | Changed Argon lamp location.
02180 | AW | AL | I020629_200304_473_B.FIT | 4 |  | I | [5 empty] | Oops, too far. Lost lines.
02180 | AW | AL | I020629_200307_356_B.FIT | 4 |  | I
02180 | AW | AL | I020629_200311_306_B.FIT | 4 |  | I
02180 | AW | AL | I020629_200553_906_B.FIT | 4 | 909 | I | -20 | 2.1 | 0 |  |  | Changed Argon lamp location.
02180 | AW | AL | I020629_200556_756_B.FIT | 4 |  | I | [5 empty] | Still no lines.
02180 | AW | AL | I020629_200559_640_B.FIT | 4 |  | I
02180 | AW | AL | I020629_200603_540_B.FIT | 4 |  | I

## 020630
YRDAY | SEQ | TYPE | Filen30es for 02181 (June 30 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Van Cleve
|  |  | Optical Bench Temp: | 128K
564449 - Collected IRFPA Image Table
02181 | AA | LVL | Levels
02181 | AA | LVL | I020630_125225_500_B.FIT | 6 | n/a | I | n/a | n/a | n/a |  |  | Dark. Reference = wood box in corner. ITOC problem. Reinitialize!
02181 | AA | LVL | I020630_125335_000_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AA | LVL | I020630_125543_700_B.FIT | 6 | 1000 | I | [5 empty] | Dark
02181 | AA | LVL | I020630_125548_133_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AA | LVL | I020630_130042_800_B.FIT | 6 | 1001 | I | [5 empty] | Dark
02181 | AA | LVL | I020630_130047_233_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | Allmodes dark
02181 | AB | DRKW | I020630_130614_433_B.FIT | 1 | 1001 | I | n/a | n/a | n/a |  |  | Dark. Temp (FPA) = ~xxxK, Temp (Optical Bench) = ~xxxK
02181 | AB | DRKW | I020630_130617_200_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130620_166_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130622_950_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130625_900_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130628_666_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130631_633_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130635_516_B.FIT | 1 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130802_450_B.FIT | 2 | 1002 | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130803_916_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130805_400_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130806_700_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130808_166_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130809_650_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130811_133_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130813_750_B.FIT | 2 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130939_683_B.FIT | 3 | 1003 | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130940_416_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130941_166_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130941_916_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130942_650_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130943_383_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130944_083_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_130946_050_B.FIT | 3 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131115_500_B.FIT | 4 | 1004 | I | [5 empty] | Dark. About 30 DN higher than 6/29, but looks OK.
02181 | AB | DRKW | I020630_131118_350_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131121_233_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131125_433_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131236_050_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131238_900_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131241_783_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131245_783_B.FIT | 4 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131414_133_B.FIT | 5 | 1005 | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131416_950_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131419_866_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131422_683_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131425_600_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131428_450_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131431_366_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131434_183_B.FIT | 5 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131602_433_B.FIT | 6 | 1006 | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131605_933_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131609_583_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131613_100_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131616_783_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131620_250_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | DRKW | I020630_131623_700_B.FIT | 6 |  | I | [5 empty] | Dark
02181 | AB | CHK | i020630_131627_133_b.FIT | 6 | [7 empty] | Dark
02181 | AC | CHK | Memory check
02181 | AC | CHK | I020630_132013_100_B.FIT | 0 | 1007 | I | n/a | n/a | n/a |  |  | Memory check mode - No data. BAD file - size is too small (only 4K).
02181 | AD | DRKW | Darks
02181 | AD | DRKW | I020630_132157_366_B.FIT | 4 | 1008 | I | [5 empty] | Test a single Mode 4. BAD file size.
02181 | AE | CHK | Memory check
02181 | AE | CHK | I020630_132503_983_B.FIT | 7 | 1009 | I | [5 empty] | First Mode 7 size looks odd. Retake two Mode 7 frames.
02181 | AE | CHK | I020630_132506_850_B.FIT | 7 |  | I
02181 | AF | UNK | Test condition not recorded
02181 | AF | UNK | I020630_132636_583_B.FIT | 4 | 1010 | I | [5 empty] | Commanded Mode 4 frame. Visually looks good. DN about the same.
02181 | AF | UNK | I020630_132751_266_B.FIT | 4 | 1011 | I | n/a | n/a | n/a |  |  | Commanded two Mode 4 frames.
02181 | AF | UNK | I020630_132754_183_B.FIT | 4 |  | I
02181 | AF | BAD | i020630_134451_083_b.FIT | [8 empty] | BAD FILE
02181 | AG | STIMIR | Stimulator
02181 | AG | STIMIR | I020630_134743_583_B.FIT | 4 | 1012 | I | [5 empty] | Stimulator ON. DAC Hex A0 (max).
02181 | AG | STIMIR | I020630_134747_683_B.FIT | 4 |  | I | [5 empty] | Took two Mode 4 frames.
02181 | AH | STIMIR | Stimulator
02181 | AH | STIMIR | I020630_135831_333_B.FIT | 1 | 224 | I | n/a | n/a | n/a |  |  | Take 8 frames for each mode (1-6) with stimulator at DAC Hex A0.
02181 | AH | STIMIR | I020630_135834_100_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135837_066_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135839_833_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135842_800_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135845_583_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135848_533_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_135852_466_B.FIT | 1 |  | I
02181 | AH | STIMIR | I020630_140019_366_B.FIT | 2 | 225 | I
02181 | AH | STIMIR | I020630_140020_833_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140022_316_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140023_616_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140025_083_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140026_566_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140028_050_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140030_666_B.FIT | 2 |  | I
02181 | AH | STIMIR | I020630_140156_550_B.FIT | 3 | 226 | I
02181 | AH | STIMIR | I020630_140157_300_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140158_033_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140158_766_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140159_516_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140200_250_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140200_983_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140202_966_B.FIT | 3 |  | I
02181 | AH | STIMIR | I020630_140332_283_B.FIT | 4 | 227 | I
02181 | AH | STIMIR | I020630_140335_116_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140338_016_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140342_016_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140452_816_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140455_650_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140458_550_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140502_300_B.FIT | 4 |  | I
02181 | AH | STIMIR | I020630_140631_700_B.FIT | 5 | 228 | I
02181 | AH | STIMIR | I020630_140634_666_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140637_433_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140640_350_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140643_166_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140645_983_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140648_933_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140651_750_B.FIT | 5 |  | I
02181 | AH | STIMIR | I020630_140820_966_B.FIT | 6 | 229 | I
02181 | AH | STIMIR | I020630_140824_483_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140827_983_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140831_483_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140834_983_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140838_516_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140842_000_B.FIT | 6 |  | I
02181 | AH | STIMIR | I020630_140845_483_B.FIT | 6 |  | I
02181 | AI | UNK | Test condition not recorded
02181 | AI | UNK | I020630_142748_983_B.FIT | 4 | 1013 | I | 10 | 1.75 |  |  |  | Take two Mode 4 images to find 500 micron slit. Got 3 pinholes. Saturated at long wavelengths.
02181 | AI | UNK | I020630_142753_183_B.FIT | 4 |  | I | [5 empty] | Used setting from 6/29. Good CO2 lines.
02181 | AI | UNK | I020630_142936_483_B.FIT | 4 | 1014 | I | 10 | 2 |  |  |  | Signal decreased.
02181 | AI | UNK | I020630_142940_633_B.FIT | 4 |  | I
02181 | AI | UNK | I020630_143105_083_B.FIT | 4 | 1015 | I | 10 | 1.5 |  |  |  | Signal decreased. Therefore, use 1.750mm for Argon test.
02181 | AI | UNK | I020630_143109_383_B.FIT | 4 |  | I
02181 | AJ | AL | Argon
02181 | AJ | AL | I020630_143645_700_B.FIT | 4 | 1016 | I | 10 | 1.75 |  |  |  | Installed Argon lamp and removed pinholes. Jaws at 2mm.
02181 | AJ | AL | I020630_143650_000_B.FIT | 4 |  | I | [5 empty] | Image saturated.
02181 | AJ | AL | I020630_144132_400_B.FIT | 4 | 1017 | I | 10 | 1.75 |  |  |  | C olumns with S/N > 10: 344, 450, 497
02181 | AJ | AL | I020630_144136_650_B.FIT | 4 |  | I
02181 | AJ | AL | Allmodes Argon center slit
02181 | AK | AL | I020630_145033_516_B.FIT | 1 | 160 | I | 10 | 1.75 |  |  |  | Take eight frames of Argon at each mode (1-6).
02181 | AK | AL | I020630_145036_283_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145039_250_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145042_016_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145044_983_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145047_750_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145050_700_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145054_650_B.FIT | 1 |  | I
02181 | AK | AL | I020630_145221_683_B.FIT | 2 | 161 | I | 10 | 1.75
02181 | AK | AL | I020630_145223_166_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145224_633_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145225_933_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145227_416_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145228_900_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145230_383_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145233_033_B.FIT | 2 |  | I
02181 | AK | AL | I020630_145357_800_B.FIT | 3 | 162 | I | 10 | 1.75
02181 | AK | AL | I020630_145358_550_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145359_283_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145400_016_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145400_766_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145401_500_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145402_216_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145404_216_B.FIT | 3 |  | I
02181 | AK | AL | I020630_145533_583_B.FIT | 4 | 163 | I | 10 | 1.75
02181 | AK | AL | I020630_145536_416_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145539_300_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145543_350_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145654_116_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145656_950_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145659_850_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145703_650_B.FIT | 4 |  | I
02181 | AK | AL | I020630_145832_050_B.FIT | 5 | 164 | I | 10 | 1.75
02181 | AK | AL | I020630_145835_016_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145837_783_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145840_750_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145843_516_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145846_483_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145849_283_B.FIT | 5 |  | I
02181 | AK | AL | I020630_145852_250_B.FIT | 5 |  | I
02181 | AK | AL | I020630_150021_450_B.FIT | 6 | 165 | I | 10 | 1.75
02181 | AK | AL | I020630_150024_950_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150028_416_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150031_933_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150035_433_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150038_966_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150042_433_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150045_766_B.FIT | 6 |  | I
02181 | AK | AL | I020630_150413_016_B.FIT | 4 | 1017 | I | 10 | 1.75 |  |  |  | Argon - Center of slit.
02181 | AK | AL | I020630_150417_316_B.FIT | 4 |  | I
02181 | AL | AL | Allmodes Argon top slit
02181 | AL | AL | I020630_150804_650_B.FIT | 1 | 160 | I | 10 | 1.75 |  |  |  | Argon - top of slit. S/N is OK.
02181 | AL | AL | I020630_150807_433_B.FIT | 1 |  | I | [5 empty] | Take eight frames of Argon at each mode (1-6).
02181 | AL | AL | I020630_150810_383_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150813_166_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150816_116_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150818_900_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150821_850_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150825_783_B.FIT | 1 |  | I
02181 | AL | AL | I020630_150952_650_B.FIT | 2 | 161 | I | 10 | 1.75
02181 | AL | AL | I020630_150954_116_B.FIT | 2 |  | I
02181 | AL | AL | I020630_150955_616_B.FIT | 2 |  | I
02181 | AL | AL | I020630_150956_900_B.FIT | 2 |  | I
02181 | AL | AL | I020630_150958_383_B.FIT | 2 |  | I
02181 | AL | AL | I020630_150959_866_B.FIT | 2 |  | I
02181 | AL | AL | I020630_151001_333_B.FIT | 2 |  | I
02181 | AL | AL | I020630_151003_950_B.FIT | 2 |  | I
02181 | AL | AL | I020630_151128_833_B.FIT | 3 | 162 | I | 10 | 1.75
02181 | AL | AL | I020630_151129_583_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151130_316_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151131_066_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151131_800_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151132_550_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151133_283_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151135_300_B.FIT | 3 |  | I
02181 | AL | AL | I020630_151304_616_B.FIT | 4 | 163 | I | 10 | 1.75
02181 | AL | AL | I020630_151307_466_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151310_350_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151314_550_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151425_116_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151427_966_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151430_850_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151434_900_B.FIT | 4 |  | I
02181 | AL | AL | I020630_151603_100_B.FIT | 5 | 164 | I | 10 | 1.75
02181 | AL | AL | I020630_151606_066_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151608_833_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151611_750_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151614_566_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151617_383_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151620_333_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151623_150_B.FIT | 5 |  | I
02181 | AL | AL | I020630_151752_433_B.FIT | 6 | 165 | I | 10 | 1.75
02181 | AL | AL | I020630_151755_983_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151759_450_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151802_916_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151806_400_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151810_000_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151813_533_B.FIT | 6 |  | I
02181 | AL | AL | I020630_151816_950_B.FIT | 6 |  | I
02181 | AM | AL | Argon bottom of slit
02181 | AM | AL | I020630_152429_600_B.FIT | 4 | 1018 | I | 10 | 1.75 |  |  |  | Argon - bottom of slit.
02181 | AM | AL | I020630_152433_950_B.FIT | 4 |  | I
02181 | AN | AL | Allmodes Argon
02181 | AN | AL | I020630_152649_800_B.FIT | 1 | 160 | I | 10 | 1.75 |  |  |  | Take eight frames of Argon at each mode (1-6).
02181 | AN | AL | I020630_152652_583_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152655_533_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152658_316_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152701_266_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152704_050_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152707_000_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152710_900_B.FIT | 1 |  | I
02181 | AN | AL | I020630_152837_850_B.FIT | 2 | 161 | I | 10 | 1.75
02181 | AN | AL | I020630_152839_333_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152840_816_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152842_116_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152843_583_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152845_066_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152846_550_B.FIT | 2 |  | I
02181 | AN | AL | I020630_152849_166_B.FIT | 2 |  | I
02181 | AN | AL | I020630_153014_033_B.FIT | 3 | 162 | I | 10 | 1.75
02181 | AN | AL | I020630_153014_766_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153015_500_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153016_250_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153016_983_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153017_716_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153018_433_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153020_433_B.FIT | 3 |  | I
02181 | AN | AL | I020630_153148_883_B.FIT | 4 | 163 | I | 10 | 1.75
02181 | AN | AL | I020630_153151_716_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153154_600_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153158_650_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153309_416_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153312_250_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153315_150_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153318_950_B.FIT | 4 |  | I
02181 | AN | AL | I020630_153448_433_B.FIT | 5 | 164 | I | 10 | 1.75
02181 | AN | AL | I020630_153451_250_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153454_166_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153456_983_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153459_900_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153502_750_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153505_566_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153508_533_B.FIT | 5 |  | I
02181 | AN | AL | I020630_153508_533_B.FIT | 6 | 165 | I | 10 | 1.75 |  |  |  | Last Argon set.
02181 | AN | AL | I020630_153636_650_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153640_216_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153643_700_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153647_200_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153650_683_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153654_166_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153657_650_B.FIT | 6 |  | I
02181 | AN | AL | I020630_153701_016_B.FIT | 6 |  | I
02181 | AO | KL | Krypton
02181 | AO | KL | I020630_154419_666_B.FIT | 4 | n/a | I | 10 | 1.75 |  |  |  | Installed Krypton lamp at center of slit.
02181 | AO | KL | I020630_154423_766_B.FIT | 4 |  | I | [5 empty] | Columns with S/N >> 10: 495, 537.
02181 | AP | KL | Allmodes Krypton
02181 | AP | KL | I020630_155146_500_B.FIT | 1 | 176 | I | 10 | 1.75 |  |  |  | Take eight frames of Krypton at each mode (1-6).
02181 | AP | KL | I020630_155149_266_B.FIT | 1 |  | I | [5 empty] | Very good signal. Krypton may be less ambiguous than Argon.
02181 | AP | KL | I020630_155152_233_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155155_000_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155157_966_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155200_750_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155203_700_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155207_633_B.FIT | 1 |  | I
02181 | AP | KL | I020630_155334_550_B.FIT | 2 | 177 | I | 10 | 1.75
02181 | AP | KL | I020630_155336_016_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155337_500_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155338_800_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155340_266_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155341_750_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155343_233_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155345_850_B.FIT | 2 |  | I
02181 | AP | KL | I020630_155511_666_B.FIT | 3 | 178 | I | 10 | 1.75
02181 | AP | KL | I020630_155512_416_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155513_150_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155513_883_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155514_633_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155515_366_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155516_100_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155518_083_B.FIT | 3 |  | I
02181 | AP | KL | I020630_155647_466_B.FIT | 4 | 179 | I | 10 | 1.75
02181 | AP | KL | I020630_155650_300_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155653_200_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155657_150_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155808_016_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155810_850_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155813_733_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155817_550_B.FIT | 4 |  | I
02181 | AP | KL | I020630_155946_900_B.FIT | 5 | 180 | I | 10 | 1.75
02181 | AP | KL | I020630_155949_816_B.FIT | 5 |  | I
02181 | AP | KL | I020630_155952_633_B.FIT | 5 |  | I
02181 | AP | KL | I020630_155955_550_B.FIT | 5 |  | I
02181 | AP | KL | I020630_155958_366_B.FIT | 5 |  | I
02181 | AP | KL | I020630_160001_133_B.FIT | 5 |  | I
02181 | AP | KL | I020630_160004_100_B.FIT | 5 |  | I
02181 | AP | KL | I020630_160006_866_B.FIT | 5 |  | I
02181 | AP | KL | I020630_160136_133_B.FIT | 6 | 181 | I | 10 | 1.75
02181 | AP | KL | I020630_160139_716_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160143_233_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160146_766_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160150_266_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160153_783_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160157_316_B.FIT | 6 |  | I
02181 | AP | KL | I020630_160200_733_B.FIT | 6 |  | I
02181 | AQ | BBJ | Blackbody (which)
02181 | AQ | BBJ | I020630_164204_216_B.FIT | 0 | n/a | I | 10 | 1.75 |  |  |  | Installed the BIG black body (BB). This is a BAD file.
02181 | AQ | BBJ | I020630_164505_433_B.FIT | 4 | 1020 | I | 10 | 1.75 |  |  |  | BB = -3.05 mV. Take two Mode 4 images. This is a test to get us onscale for tomorrow's tests.
02181 | AQ | BBJ | I020630_164509_583_B.FIT | 4 |  | I | [5 empty] | Good H2O lines.
02181 | AQ | BBJ | I020630_164944_900_B.FIT | 4 | 1021 | I | 10 | 1.75 |  |  |  | BB = -3.72 mV (~450 deg C). 100 deg C = ~1 mV.
02181 | AQ | BBJ | I020630_164949_000_B.FIT | 4 |  | I | [5 empty] | Saturated down to pixel 650. 1000 DN occurs around pixel 400.
02181 | AQ | BBJ | I020630_170017_400_B.FIT | 4 | 1022 | I | 10 | 1.75 |  |  |  | BB = -4.00 mV (~500 deg C). BEST SETTLED.
02181 | AQ | BBJ | I020630_170021_700_B.FIT | 4 |  | I | [5 empty] | 1000 DN occurs around pixel 300.
02181 | AQ | BBJ | I020630_170928_666_B.FIT | 4 | 1023 | I | 10 | 1.75 |  |  |  | Heat BB = -5.0 mV.
02181 | AQ | BBJ | I020630_170932_966_B.FIT | 4 |  | I
02181 | AQ | BBJ | I020630_171812_350_B.FIT | 4 | 1024 | I | 10 | 1.75 |  |  |  | Heat BB = -6.2 mV (693 dec C).
02181 | AQ | BBJ | I020630_171816_450_B.FIT | 4 |  | I | [5 empty] | 1000 DN occurs around at shortest wavelength. This is NOT a typo!
02181 | AQ | BBJ | I020630_172652_400_B.FIT | 4 | 1025 | I | 10 | 1.75 |  |  |  | Heat BB = -7.2 mV (~800 deg C).
02181 | AQ | BBJ | I020630_172656_600_B.FIT | 4 |  | I
02181 | AQ | BBJ | I020630_173406_883_B.FIT | 4 | 1026 | I | 10 | 1.75 |  |  |  | Heat BB = -7.7 mV (~900 deg C).
02181 | AQ | BBJ | I020630_173411_083_B.FIT | 4 |  | I | [5 empty] | 2000 DN occurs at shortest wavelength. Turned off BB.
02181 | AQ | BBJ | I020630_174428_583_B.FIT | 0 | n/a | I | n/a | n/a | n/a |  |  | BAD file isize; 1-byte image.

## 020701
YRDAY | SEQ | TYPE | Filenames for 02182 (July 1 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Smith (AM) / Van Cleve (PM)
|  |  | Optical Bench Temp: | 135K
564449 - Collected IRFPA Image Table
02182 | AA | BAD | Bad file
02182 | AA | BAD | I020701_013739_933_B.FIT | 0 | n/a | I | n/a | n/a | n/a |  |  | n/a
02182 | AB | LVL | Level check
02182 | AB | LVL | I020701_100408_933_B.FIT | 6 | 1100 | I | n/a | n/a | n/a |  |  | n/a
02182 | AB | LVL | I020701_100413_366_B.FIT | 6 |  | I | [5 empty] | n/a
02182 | AC | DRKW | Dark. Reference = wood box in corner.
02182 | AC | DRKW | I020701_100819_933_B.FIT | 1 | 1101 | I | n/a | n/a | n/a |  |  | Dark. Reference = wood box in corner.
02182 | AC | DRKW | I020701_100822_716_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100825_683_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100828_450_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100831_416_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100834_183_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100837_150_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_100841_033_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101009_100_B.FIT | 2 |  | I | n/a | n/a | n/a |  |  | Dark
02182 | AC | DRKW | I020701_101010_566_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101012_050_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101013_350_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101014_833_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101016_316_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101017_783_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101020_400_B.FIT | 2 |  | I |  | n/a |  |  |  | Dark
02182 | AC | DRKW | I020701_101146_316_B.FIT | 3 |  | I | n/a |  | n/a |  |  | Dark
02182 | AC | DRKW | I020701_101147_066_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101147_800_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101148_533_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101149_283_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101150_016_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101150_750_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101152_733_B.FIT | 3 |  | I |  | n/a |  |  |  | Dark
02182 | AC | DRKW | I020701_101322_116_B.FIT | 4 | 1 | I | n/a |  | n/a |  |  | Dark
02182 | AC | DRKW | I020701_101324_966_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101327_850_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101331_850_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101442_683_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101445_533_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101448_416_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101452_316_B.FIT | 4 |  | I |  | n/a |  |  |  | Dark
02182 | AC | DRKW | I020701_101621_583_B.FIT | 5 | 2 | I | n/a |  | n/a |  |  | Dark
02182 | AC | DRKW | I020701_101624_400_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101627_366_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101630_083_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101633_050_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101635_866_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101638_783_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101641_633_B.FIT | 5 |  | I |  | n/a |  |  |  | Dark
02182 | AC | DRKW | I020701_101810_783_B.FIT | 6 | 3 | I | n/a |  | n/a |  |  | Dark
02182 | AC | DRKW | I020701_101814_300_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101817_816_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101821_416_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101825_016_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101828_633_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101832_150_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AC | DRKW | I020701_101835_450_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AD | CHK | Memory check mode. |  |  |  |  | n/a
02182 | AD | CHK | I020701_102144_316_B.FIT | 7 | 4 | I | n/a |  | n/a |  |  | Memory check mode.
02182 | AD | CHK | I020701_102147_183_B.FIT | 7 |  | I | [5 empty] | Memory check mode.
02182 | AE | MSK | Knife Edge |  |  |  |  | 1.7
02182 | AE | MSK | I020701_102527_650_B.FIT | 4 | 1108 | I | 9 | n/a |  |  |  | Begin fine knife-edge tests at -3mm, -1mm, +1mm. D. Gallagher.
02182 | AE | MSK | I020701_102531_800_B.FIT | 0 | n/a | I | n/a | 1.7 |  |  |  | Bad frame.
02182 | AE | MSK | I020701_103917_733_B.FIT | 4 | 1109 | I | 9 | 1.6
02182 | AE | MSK | I020701_104251_216_B.FIT | 4 | 1110 | I | 9 | 1.8
02182 | AE | MSK | I020701_104429_833_B.FIT | 4 | 1111 | I | 9 | 1.9
02182 | AE | MSK | I020701_104616_100_B.FIT | 4 | 1112 | I | 9 | 2
02182 | AE | MSK | I020701_104721_650_B.FIT | 4 | 1113 | I | 9 | 2.1
02182 | AE | MSK | I020701_104808_783_B.FIT | 4 | 1114 | I | 9 | 1.5
02182 | AE | MSK | I020701_104916_100_B.FIT | 4 | 1115 | I | 9 | 1.4
02182 | AE | MSK | I020701_104956_583_B.FIT | 4 | 1116 | I | 9 | 1.3
02182 | AE | MSK | I020701_105042_783_B.FIT | 4 | 1117 | I | 9 | 1.75
02182 | AE | MSK | I020701_105440_316_B.FIT | 4 | 1118 | I | 11 | 1.75
02182 | AE | MSK | I020701_105538_450_B.FIT | 4 | 1119 | I | 13 | 1.75
02182 | AE | MSK | I020701_105633_650_B.FIT | 4 | 1120 | I | 15 | 1.75
02182 | AE | MSK | I020701_105720_550_B.FIT | 4 | 1121 | I | 17 | 1.75
02182 | AE | MSK | I020701_105810_083_B.FIT | 4 | 1122 | I | 19 | 1.8
02182 | AE | MSK | I020701_105927_783_B.FIT | 4 | 1123 | I | 15 | 1.9
02182 | AE | MSK | I020701_110008_150_B.FIT | 4 | 1124 | I | 15 | 2
02182 | AE | MSK | I020701_110048_800_B.FIT | 4 | 1125 | I | 15 | 2.1
02182 | AE | MSK | I020701_110132_050_B.FIT | 4 | 1126 | I | 15 | 2.2
02182 | AE | MSK | I020701_110232_883_B.FIT | 4 | 1127 | I | 15 | 1.6
02182 | AE | MSK | I020701_110406_750_B.FIT | 4 | 1128 | I | 15 | 1.5
02182 | AE | MSK | I020701_110444_633_B.FIT | 4 | 1129 | I | 15 | 1.4
02182 | AE | MSK | I020701_110523_566_B.FIT | 4 | 1130 | I | 15 | 1.3
02182 | AE | MSK | I020701_110600_900_B.FIT | 4 | 1131 | I | 15 | 1.75
02182 | AE | MSK | I020701_110808_633_B.FIT | 4 | 1132 | I | 7 | 1.75
02182 | AE | MSK | I020701_110856_050_B.FIT | 4 | 1133 | I | 5 | 1.75
02182 | AE | MSK | I020701_110942_800_B.FIT | 4 | 1134 | I | 3 | 1.75
02182 | AE | MSK | I020701_111024_333_B.FIT | 4 | 1135 | I | 1 | 1.75
02182 | AE | MSK | I020701_111111_450_B.FIT | 4 | 1136 | I | -1 | 1.75
02182 | AE | MSK | I020701_111201_416_B.FIT | 4 | 1137 | I | -3 | 1.75
02182 | AE | MSK | I020701_111239_333_B.FIT | 4 | 1138 | I | -5 | 1.75
02182 | AE | MSK | I020701_111326_566_B.FIT | 4 | 1139 | I | -7 | 1.75
02182 | AE | MSK | I020701_111437_033_B.FIT | 4 | 1140 | I | -9
02182 | AF | MSK | Change Z-Focus to -1mm. |  |  |  |  | 1.75
02182 | AF | MSK | I020701_111536_016_B.FIT | 4 | 1141 | I | -1 | 1.7 |  |  |  | Change Z-Focus to -1mm. Best focus seems to be between -3mm and +1mm. Take frames at -3mm, -1mm, and +1mm.
02182 | AF | MSK | I020701_112550_366_B.FIT | 4 | 1142 | I | -1 | 1.65
02182 | AF | MSK | I020701_112650_950_B.FIT | 4 | 1143 | I | -1 | 1.6
02182 | AF | MSK | I020701_112726_466_B.FIT | 4 | 1144 | I | -1 | 1.55
02182 | AF | MSK | I020701_112803_116_B.FIT | 4 | 1145 | I | -1 | 1.5
02182 | AF | MSK | I020701_112854_716_B.FIT | 4 | 1146 | I | -1 | 1.45
02182 | AF | MSK | I020701_112945_266_B.FIT | 4 | 1147 | I | -1 | 1.4
02182 | AF | MSK | I020701_113022_050_B.FIT | 4 | 1148 | I | -1 | 1.8
02182 | AF | MSK | I020701_113103_000_B.FIT | 4 | 1149 | I | -1 | 1.85
02182 | AF | MSK | I020701_113205_616_B.FIT | 4 | 1150 | I | -1 | 1.9
02182 | AF | MSK | I020701_113250_466_B.FIT | 4 | 1151 | I | -1 | 1.95
02182 | AF | MSK | I020701_113335_616_B.FIT | 4 | 1152 | I | -1 | 2
02182 | AF | MSK | I020701_113415_150_B.FIT | 4 | 1153 | I | -1 | 1.7
02182 | AF | MSK | I020701_114220_450_B.FIT | 4 | 1154 | I | -3 | 1.65 |  |  |  | Change Z-Focus to -3mm.
02182 | AF | MSK | I020701_114307_166_B.FIT | 4 | 1155 | I | -3 | 1.6
02182 | AF | MSK | I020701_114350_350_B.FIT | 4 | 1156 | I | -3 | 1.55
02182 | AF | MSK | I020701_114514_766_B.FIT | 4 | 1157 | I | -3 | 1.5
02182 | AF | MSK | I020701_114600_316_B.FIT | 4 | 1158 | I | -3 | 1.45
02182 | AF | MSK | I020701_114640_800_B.FIT | 4 | 1159 | I | -3 | 1.8
02182 | AF | MSK | I020701_114743_433_B.FIT | 4 | 1160 | I | -3 | 1.85
02182 | AF | MSK | I020701_114822_916_B.FIT | 4 | 1161 | I | -3 | 1.9
02182 | AF | MSK | I020701_114907_300_B.FIT | 4 | 1162 | I | -3 | 1.95
02182 | AF | MSK | I020701_114942_616_B.FIT | 4 | 1163 | I | -3 | 2
02182 | AF | MSK | I020701_115019_166_B.FIT | 4 | 1164 | I | -3 | 1.7
02182 | AF | MSK | I020701_115114_516_B.FIT | 4 | 1165 | I | 1 | 1.65 |  |  |  | Change Z-Focus to 1mm.
02182 | AF | MSK | I020701_115156_916_B.FIT | 4 | 1166 | I | 1 | 1.6
02182 | AF | MSK | I020701_115234_100_B.FIT | 4 | 1167 | I | 1 | 1.55
02182 | AF | MSK | I020701_115310_050_B.FIT | 4 | 1168 | I | 1 | 1.5
02182 | AF | MSK | I020701_115348_333_B.FIT | 4 | 1169 | I | 1 | 1.45
02182 | AF | MSK | I020701_115435_500_B.FIT | 4 | 1170 | I | 1 | 1.8
02182 | AF | MSK | I020701_115520_550_B.FIT | 4 | 1171 | I | 1 | 1.85
02182 | AF | MSK | I020701_115600_216_B.FIT | 4 | 1172 | I | 1 | 1.9
02182 | AF | MSK | I020701_115703_000_B.FIT | 4 | 1173 | I | 1 | 1.95
02182 | AF | MSK | I020701_115742_266_B.FIT | 4 | 1174 | I | 1 | 2
02182 | AF | MSK | I020701_115824_266_B.FIT | 4 | 1175 | I | 1 | 1.75
02182 | AF | MSK | I020701_123906_516_B.FIT | 4 | 1176 | I | 9 | 1.75 | 0
02182 | AF | MSK | I020701_124133_883_B.FIT | 4 | 1177 | I | 9 | 1.75 | -2.5
02182 | AF | MSK | I020701_124233_050_B.FIT | 4 | 1178 | I | 9 | 1.75 | -5
02182 | AF | MSK | I020701_124339_550_B.FIT | 4 | 1179 | I | 9 | 1.75 | -7.5
02182 | AF | MSK | I020701_124432_600_B.FIT | 4 | 1180 | I | 9 | 1.75 | -10
02182 | AF | MSK | I020701_124726_716_B.FIT | 4 | 1181 | I | 9 |  | 0 |  |  | Block bottom pinhole.
02182 | AG | BBJ | Black Body=2250 |  |  |  |  | 1.75
02182 | AG | BBJ | I020701_131133_833_B.FIT | 4 | 1182 | I | 4 |  |  |  |  | Installed the big BB. BB Temp = -1.656 mV. Aperature open at 1.00. BB set point = 2250. Used dark I020701_101322_116_B.FIT.
02182 | AG | BBJ | I020701_131137_833_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_132005_800_B.FIT | 4 | 1183 | I | 4 |  |  |  |  | BB Temp = -1.656 mV. Closed aperature.
02182 | AG | BBJ | I020701_132009_850_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_132206_083_B.FIT | 4 | 1184 | I | 4 |  |  |  |  | Bad frame.
02182 | AG | BBJ | I020701_132210_283_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_132332_200_B.FIT | 4 | 1185 | I | 4 |  |  |  |  | BB Temp = -1.656 mV. Closed aperature with Mylar.
02182 | AG | BBJ | I020701_132336_600_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_135713_250_B.FIT | 4 | 1186 | I | 4 |  |  |  |  | BB Temp = -2.03 mV. Closed aperature BB's mask. BB set point = 2750.
02182 | AG | BBJ | I020701_135717_450_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_135905_466_B.FIT | 4 | 1187 | I | 4 |  |  |  |  | BB Temp = -2.02 mV. Aperature open.
02182 | AG | BBJ | I020701_135909_766_B.FIT | 4 |  | I |  | 1.75
02182 | AG | BBJ | I020701_141415_833_B.FIT | 4 | 1188 | I | 4 |  |  |  |  | Collect four Mode 4. BB Temp = -2.417 mV. BB set point = 2750.
02182 | AG | BBJ | I020701_141418_666_B.FIT | 4 |  | I
02182 | AG | BBJ | I020701_141421_566_B.FIT | 4 |  | I
02182 | AG | BBJ | I020701_141425_516_B.FIT | 4 |  | I
02182 | AH | DRKM | BB dark with mylar |  |  |  |  | 1.75
02182 | AH | DRKM | I020701_142153_816_B.FIT | 1 | 70 | I | 4 |  |  |  |  | Closed aperature with Mylar and took darks (gives same path length as BB data frames).
02182 | AH | DRKM | I020701_142156_583_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142159_550_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142202_333_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142205_283_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142208_050_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142211_016_B.FIT | 1 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142214_900_B.FIT | 1 |  | I |  | 1.75 |  |  |  | Dark
02182 | AH | DRKM | I020701_142340_950_B.FIT | 2 | 71 | I | 4 |  |  |  |  | Dark
02182 | AH | DRKM | I020701_142342_450_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142343_916_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142345_216_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142346_700_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142348_166_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142349_650_B.FIT | 2 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142352_216_B.FIT | 2 |  | I |  | 1.75 |  |  |  | Dark
02182 | AH | DRKM | I020701_142517_100_B.FIT | 3 | 72 | I | 4 |  |  |  |  | Dark
02182 | AH | DRKM | I020701_142517_833_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142518_583_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142519_316_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142520_050_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142520_800_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142521_533_B.FIT | 3 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142523_550_B.FIT | 3 |  | I |  | 1.75 |  |  |  | Dark
02182 | AH | DRKM | I020701_142651_700_B.FIT | 4 | 73 | I | 4 |  |  |  |  | Dark
02182 | AH | DRKM | I020701_142654_550_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142657_433_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142701_583_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142812_133_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142814_966_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142817_850_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142821_950_B.FIT | 4 |  | I |  | 1.75 |  |  |  | Dark
02182 | AH | DRKM | I020701_142950_600_B.FIT | 5 | 74 | I | 4 |  |  |  |  | Dark
02182 | AH | DRKM | I020701_142953_566_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142956_416_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_142959_333_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143002_150_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143005_066_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143007_883_B.FIT | 5 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143010_650_B.FIT | 5 |  | I |  | 1.75 |  |  |  | Dark
02182 | AH | DRKM | I020701_143139_600_B.FIT | 6 | 75 | I | 4 |  |  |  |  | Dark
02182 | AH | DRKM | I020701_143143_033_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143146_533_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143150_100_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143153_616_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143157_133_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143200_616_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AH | DRKM | I020701_143204_050_B.FIT | 6 |  | I | [5 empty] | Dark
02182 | AI | BBJ | BBJ=? |  |  |  |  | 1.75
02182 | AI | BBJ | I020701_143408_366_B.FIT | 1 | 76 | I | 4 |  |  |  |  | Removed Mylar mask. Aperature open at 1.00. BB Temp = -2.4315 mV (equilabrated).
02182 | AI | BBJ | I020701_143411_150_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143414_100_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143416_883_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143419_833_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143422_616_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143425_566_B.FIT | 1 |  | I
02182 | AI | BBJ | I020701_143429_466_B.FIT | 1 |  | I |  | 1.75
02182 | AI | BBJ | I020701_143556_333_B.FIT | 2 | 77 | I | 4 |  |  |  |  | Big BB
02182 | AI | BBJ | I020701_143557_816_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143559_300_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143600_600_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143602_083_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143603_550_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143605_033_B.FIT | 2 |  | I
02182 | AI | BBJ | I020701_143607_600_B.FIT | 2 |  | I |  | 1.75
02182 | AI | BBJ | I020701_143733_533_B.FIT | 3 | 78 | I | 4 |  |  |  |  | Big BB
02182 | AI | BBJ | I020701_143734_266_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143735_016_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143735_750_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143736_483_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143737_233_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143737_966_B.FIT | 3 |  | I
02182 | AI | BBJ | I020701_143739_983_B.FIT | 3 |  | I |  | 1.75
02182 | AI | BBJ | I020701_143909_300_B.FIT | 4 | 79 | I | 4 |  |  |  |  | Big BB
02182 | AI | BBJ | I020701_143912_183_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_143915_016_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_143919_083_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_144029_833_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_144032_666_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_144035_566_B.FIT | 4 |  | I
02182 | AI | BBJ | I020701_144039_566_B.FIT | 4 |  | I |  | 1.75
02182 | AI | BBJ | I020701_144208_650_B.FIT | 5 | 80 | I | 4 |  |  |  |  | Big BB
02182 | AI | BBJ | I020701_144211_566_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144214_383_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144217_300_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144220_150_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144223_016_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144225_883_B.FIT | 5 |  | I
02182 | AI | BBJ | I020701_144228_700_B.FIT | 5 |  | I |  | 1.75
02182 | AI | BBJ | I020701_144357_950_B.FIT | 6 | 81 | I | 4 |  |  |  |  | BB Temp = -2.4337 mV
02182 | AI | BBJ | I020701_144401_400_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144404_900_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144408_500_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144411_983_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144415_566_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144419_066_B.FIT | 6 |  | I
02182 | AI | BBJ | I020701_144422_466_B.FIT | 6 |  | I
02182 | AJ | BBJ | BB fill tests |  |  |  |  | 1.75
02182 | AJ | BBJ | I020701_145018_016_B.FIT | 4 | 1189 | I | 4 |  |  |  |  | Moved BB 0.25-inch up (towards roof). BB Temp = -2.430 mV (~325 deg C). Test translation of BB effects on flat field at BB Temp ~325 deg C. Use 1426511 as dark. Get a slightly off-center gaussian-like plot for a column slice (around x = 790).
02182 | AJ | BBJ | I020701_145022_116_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_145655_716_B.FIT | 4 | 1190 | I | 4 |  |  |  |  | Moved BB down as far as it will go (to -0.125 inch). BB Temp = -2.424 mV. This motion lost the signal (down to ~250 DN max).
02182 | AJ | BBJ | I020701_145659_766_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_151014_700_B.FIT | 4 | 1191 | I | 4 |  |  |  |  | Moved BB to center of the flip mirror and in (to the edge of the optical bench). Produced a "flatter" column slice (around x=800). Image even but signal lower.
02182 | AJ | BBJ | I020701_151018_850_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_152110_000_B.FIT | 4 | 1192 | I | 4 |  |  |  |  | Moved BB position. BB Temp = -2.425 mV. Slit seems to be filled but signal is very low (no saturation). Is BB not in slit?
02182 | AJ | BBJ | I020701_152114_200_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_152727_583_B.FIT | 4 | 1193 | I | 4 |  |  |  |  | BB still at edge of table. Moved BB up to the ~starting height. and adjusted mirror. BB Temp = - 2.418 mV. GOOD frame! Got a nearly flat, column slice near pix 623-632 and DN ~1500, but fall off from y-pix 0-100.
02182 | AJ | BBJ | I020701_152731_783_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_153251_883_B.FIT | 4 | 1194 | I | 4 |  |  |  |  | Try another veritcal move. Move BB down by one turn. BB Temp = -2.408. Image moved down as expected. And ~flat at y = 100-300. But signal fell off from 0-100. Instrumental fall off at bottom?
02182 | AJ | BBJ | I020701_153256_083_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_153529_366_B.FIT | 4 | 1195 | I | 4 |  |  |  |  | Moved BB uo one turn (back to 152727 location). Again, another dropoff at y = 0-100.
02182 | AJ | BBJ | I020701_153533_366_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_153753_450_B.FIT | 4 | 1196 | I | 4 |  |  |  |  | Move BB up 1/2 turn. Ok, we have good SN coverage of most of the array. Again, another dropoff at y = 0-100! Seems to imply an instrumental dropoff at bottom.
02182 | AJ | BBJ | I020701_153757_850_B.FIT | 4 |  | I |  | 1.75
02182 | AJ | BBJ | I020701_154158_333_B.FIT | 4 | 1197 | I | 4 |  |  |  |  | Moved BB ~ 0.5-inch to the left (towards the vac chamber). Signal strength, image, and coluimn slice plot ~identical to 153757. Some horizontal change has little effect.
02182 | AJ | BBJ | I020701_154202_483_B.FIT | 4 |  | I
02182 | AK | AL | Argon lamp |  |  |  |  | 1.75
02182 | AK | AL | I020701_155439_266_B.FIT | 4 | 1198 | I | 4 |  |  |  |  | Installed Argon lamp with safety glass. Slit in. The safety glass causes the line to spread dramatically in the spatial direction (when compared to 06/30 spectra). Good spectral smile. Brightest lines at x = 340 and 492. SN ~10, close to successful test criteria. Move to Kryton. Use dark from BB test, 142651.
02182 | AK | AL | I020701_155443_466_B.FIT | 4 |  | I
02182 | AL | KL | Krypton |  |  |  |  | 1.75
02182 | AL | KL | I020701_160633_566_B.FIT | 4 | 1199 | I | 4 |  |  |  |  | Installed Krypton lamp without the safety glass. Slit in. We noticed the U-shape of the bulb in the lines. Brightest lines at x = 491 and 532. SN ~10. Image shifted about 4-5 pixels from yesterday. SIM Temp ~135K. Use dark from BB test, 142651.
02182 | AL | KL | I020701_160637_666_B.FIT | 4 |  | I |  | 1.75
02182 | AL | KL | I020701_161145_883_B.FIT | 4 | 1200 | I | 4 |  |  |  |  | Krypton lamp with safety glass. Lines lengthened and filled most of the spatial direction Use dark from BB test, 142651.
02182 | AL | KL | I020701_161149_983_B.FIT | 4 |  | I
02182 | AM | TLG | Tungsten with gold IP |  |  |  |  | 1.75
02182 | AM | TLG | I020701_164855_416_B.FIT | 4 | 1201 | I | 4 |  |  |  |  | Installed gold IR integrating sphere with Tungsten lamp at 15 Amp. Slit removed. Only top 2/3 of image illuminated. Can see top edge of detector in most of this image. Need more signal; Move sphere about 100mm down. Use dark from BB test, 142651.
02182 | AM | TLG | I020701_164859_516_B.FIT | 4 |  | I |  | 1.75
02182 | AM | TLG | I020701_165745_000_B.FIT | 4 | 1202 | I | 4 |  |  |  |  | Moved IR sphere down about 1cm. Got a good vertical center. DN ~ 100 at y=430. DN ~800 at y=650.
02182 | AM | TLG | I020701_165749_200_B.FIT | 4 |  | I
02182 | AN | BBJ | BB mica test |  |  |  |  | 1.75
02182 | AN | BBJ | I020701_170934_600_B.FIT | 4 | 1203 | I | 4 |  |  |  |  | Mica film. Held film in front of BB exit. Big BB at 2.03 K. Use dark from BB test, 142651.
02182 | AN | BBJ | I020701_170939_000_B.FIT | 4 |  | I
02182 | AO | BBJ | BB Polystyrene |  |  |  |  | 1.75
02182 | AO | BBJ | I020701_171137_200_B.FIT | 4 | 1204 | I | 4 |  |  |  |  | Polystyrene film. Held film in front of BB exit. Big BB at 2.02 K. Saw absorption around 3.3 micron (~x=800) as expected. Polystyrene caused the absorption in 06-30 STIM images (eg, 134743). Use dark from BB test, 142651.
02182 | AO | BBJ | I020701_171141_300_B.FIT | 4 |  | I
02182 | AP | STIMIR | Internal stimulus |  |  |  |  | 1.75
02182 | AP | STIMIR | I020701_172628_800_B.FIT | 4 | 1205 | I | 4 |  |  |  |  | STIM flash while waiting for SIRTF BB to warm up. Slit removed. DAC Hex A0 (max setting).
02182 | AP | STIMIR | I020701_172631_633_B.FIT | 4 |  | I
02182 | AP | STIMIR | I020701_172634_533_B.FIT | 4 |  | I | [5 empty] | STIM off (DAC Hex 00).
02182 | AP | STIMIR | I020701_172638_333_B.FIT | 4 |  | I
02182 | AQ | DRKW | SIRTF BB=300 |  |  |  |  | 1.75
02182 | AQ | DRKW | I020701_173139_766_B.FIT | 4 | 1206 | I | 4 |  |  |  |  | SIRTF BB at 300 deg F (150 deg C). Oops - Box in the way. This is a GOOD DARK for the next few tests!
02182 | AR | BBS | SIRTF BB=300 to 450
02182 | AR | BBS | I020701_173143_833_B.FIT | 4 |  | I |  | 1.75
02182 | AR | BBS | I020701_173312_833_B.FIT | 4 | 1207 | I | 4 |  |  |  |  | SIRTF BB at 300 deg F. GOOD FLAT! Max ~1800 DN and slit is filled and image (a column slit) is flat! Did Big BB show vignetting off edge of mirror?
02182 | AR | BBS | I020701_173317_033_B.FIT | 4 |  | I |  | 1.75
02182 | AR | BBS | I020701_174350_350_B.FIT | 4 | 1208 | I | 4 |  |  |  |  | SIRTF BB at 450 deg F (230 deg C). Low signal iat blue (60-100 DN). DN = 1000 at x = 700.
02182 | AR | BBS | I020701_174354_450_B.FIT | 4 |  | I
02182 | AS | BBS | Sirtf BB with Polystyrene |  |  |  |  | 1.75
02182 | AS | BBS | I020701_175027_133_B.FIT | 4 | 1209 | I | 4 |  |  |  |  | Polystyrene film with SIRTF BB as it cools (< 230 dec C). Try to get lines further from saturation to the right. See absorption around 3.3 micron (~x=800) as expected and further from saturation.
02182 | AS | BBS | I020701_175031_333_B.FIT | 4 |  | I
02182 | AT | TLW | Tungsten lamp, white integrating sphere |  |  |  |  | 1.75
02182 | AT | TLW | I020701_182802_316_B.FIT | 4 | 1210 | I | 4 |  |  |  |  | Installed blue integrating sphere (uncalibrated) with Tungsten lamp at 15 Amp. Slit in. Centered on slit. Highest DN ~100 at x=600+/-100 pix. SN ~5. Column slice at x=626-635 is flat.
02182 | AT | TLW | I020701_182806_566_B.FIT | 4 |  | I
02182 | AU | DRKW | Darks |  |  |  |  | 1.75
02182 | AU | DRKW | I020701_183255_933_B.FIT | 4 | 1211 | I | 4 |  |  |  |  | Darks to be applied to the data frames from the blue integrating sphere (above).
02182 | AU | DRKW | I020701_183258_766_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183301_650_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183304_500_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183307_383_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183310_216_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183313_116_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AU | DRKW | I020701_183316_916_B.FIT | 4 |  | I | [5 empty] | Dark
02182 | AV | TL12 | Tungsten lamp, latent light |  |  |  |  | 1.75
02182 | AV | TL12 | I020701_184336_050_B.FIT | 4 | 1212 | I | 4 |  |  |  |  | Latent light test with 2mm slit horizontal. Tungsten lamp at 12 Amp. Take a context image in Mode 4. Use dark 142651 (darks used for big BB test).
02182 | AV | TL12 | I020701_184340_300_B.FIT | 4 |  | I
02182 | AW | TL12 | Tungsten lamp, latent light |  |  |  |  | 1.75
02182 | AW | TL12 | I020701_184817_716_B.FIT | 6 | 1213 | I | 4 |  |  |  |  | Latent light test in Mode 6. Take 16 images. Masked at 8 sec and unmasked at 18 sec. Difference each pair of images via to quickly see results. Bummer, DIVe is cumbersome with this mode (images come in 2 frames that must be split). Decided to leave mode 6 analysis for later. For further analysis, use dark 143139 of the BB darks sequence.
02182 | AW | TL12 | I020701_184821_166_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184824_650_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184828_233_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184831_733_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184835_333_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184838_816_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184842_366_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184845_916_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184849_483_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184852_983_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184856_433_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184900_016_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184903_533_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184907_066_B.FIT | 6 |  | I
02182 | AW | TL12 | I020701_184910_466_B.FIT | 6 |  | I
02182 | AX | TL12 | Tungsten lamp, latent light |  |  |  |  | 1.75
02182 | AX | TL12 | I020701_185933_466_B.FIT | 4 | 1214 | I | 4 |  |  |  |  | Latent light test in Mode 4 because DIVe can easily do a running difference. Take 16 images. Mask at 6 sec and unmask at 14 sec. NOTE: DIVe crashes when we try a running, pair difference. Resorted to simple pair differences. We masked at 6 seconds and removed at 14. For further analysis, use dark 142651 of the BB darks sequence.
02182 | AX | TL12 | I020701_185936_300_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_185939_200_B.FIT | 4 |  | I | [5 empty] | Got some signal. Box in motion.
02182 | AX | TL12 | I020701_185942_033_B.FIT | 4 |  | I | [5 empty] | Losing signal. Latent image?
02182 | AX | TL12 | I020701_185944_916_B.FIT | 4 |  | I | [5 empty] | Box being removed? Signal increasing. WANT at least 6 frames to be dark for this test.
02182 | AX | TL12 | I020701_185947_766_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_185950_650_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_185953_483_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_185956_383_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_185959_216_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190002_100_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190004_950_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190007_833_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190010_666_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190013_516_B.FIT | 4 |  | I | [5 empty] | Got good signal.
02182 | AX | TL12 | I020701_190017_000_B.FIT | 4 |  | I |  | 1.75 |  |  |  | Got good signal.
02182 | AX | TL12 | I020701_191109_583_B.FIT | 4 | 1215 | I | 4 |  |  |  |  | Latent light test in Mode 4. Masked at 8 sec and unmasked at 36 sec. For further analysis, use dark 142651 of the BB darks sequence.
02182 | AX | TL12 | I020701_191112_433_B.FIT | 4 |  | I | [5 empty] | Signal.
02182 | AX | TL12 | I020701_191115_316_B.FIT | 4 |  | I | [5 empty] | Signal decreasing.
02182 | AX | TL12 | I020701_191118_166_B.FIT | 4 |  | I | [5 empty] | Ghost.
02182 | AX | TL12 | I020701_191121_050_B.FIT | 4 |  | I | [5 empty] | Tiny bit of ghost.
02182 | AX | TL12 | I020701_191123_883_B.FIT | 4 |  | I | [5 empty] | Less.
02182 | AX | TL12 | I020701_191126_766_B.FIT | 4 |  | I | [5 empty] | Noise between 6 & 7
02182 | AX | TL12 | I020701_191129_616_B.FIT | 4 |  | I | [5 empty] | Noise between 7 & 8
02182 | AX | TL12 | I020701_191132_500_B.FIT | 4 |  | I | [5 empty] | Noise between 8 & 9
02182 | AX | TL12 | I020701_191135_350_B.FIT | 4 |  | I | [5 empty] | Noise between 9 & 10
02182 | AX | TL12 | I020701_191138_233_B.FIT | 4 |  | I | [5 empty] | Noise between 10 & 11
02182 | AX | TL12 | I020701_191141_066_B.FIT | 4 |  | I | [5 empty] | Noise between 11 & 12
02182 | AX | TL12 | I020701_191143_916_B.FIT | 4 |  | I | [5 empty] | Some signal.
02182 | AX | TL12 | I020701_191146_800_B.FIT | 4 |  | I | [5 empty] | Partial signal.
02182 | AX | TL12 | I020701_191149_650_B.FIT | 4 |  | I | [5 empty] | Full signal.
02182 | AX | TL12 | I020701_191152_983_B.FIT | 4 |  | I
02182 | AY | TL12 | Tungsten lamp, latent light |  |  |  |  | 1.75
02182 | AY | TL12 | I020701_191805_366_B.FIT | 6 | 1215 | I | 4 |  |  |  |  | Latent light test in Mode 6. Masked at 10 sec and unmasked at 36 sec. Analyze later. For further analysis, use dark 143139 of the BB darks sequence.
02182 | AY | TL12 | I020701_191808_850_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191812_300_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191815_800_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191819_333_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191822_833_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191826_383_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191829_900_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191833_366_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191836_916_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191840_366_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191843_833_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191847_383_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191850_916_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191854_333_B.FIT | 6 |  | I
02182 | AY | TL12 | I020701_191857_850_B.FIT | 6 |  | I

## 020702
YRDAY | SEQ | TYPE | Filenames for 02178 (June 27 2002) | Mode | Exp. ID | V or I | Focus (mm) | X-slit (mm) | Vert (mm) | Tdelay (ms) | Filter (pos) | Comments
|  |  | Operator: | Van Cleve
|  |  | Optical Bench Temp: | 141K
564449 - Collected IRFPA Image Table
02183 | AA | LVL | Level check
02183 | AA | LVL | I020702_101016_700_B.FIT | 6 | 1300 | I | n/a | n/a | n/a |  |  | n/a
02183 | AA | LVL | I020702_101020_216_B.FIT | 6 |  | I | [5 empty] | n/a
02183 | AA | LVL | Level check
02183 | AA | LVL | I020702_101332_550_B.FIT | 0 | n/a | I | n/a | n/a | n/a |  |  | n/a
02183 | AB | DRKW | Wood-box Darks. | [8 empty] | Wood-box Darks.
02183 | AB | DRKW | I020702_101543_350_B.FIT | 1 | 1301 | I | n/a | n/a | n/a |  |  | Dark. Reference = wood box in corner.
02183 | AB | DRKW | I020702_101546_133_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101549_083_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101551_866_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101554_816_B.FIT | 1 | 1 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101557_583_B.FIT | 1 | 1302 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101600_550_B.FIT | 1 | 1 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101604_450_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101731_383_B.FIT | 2 | 1 | I | n/a | n/a | n/a |  |  | Dark
02183 | AB | DRKW | I020702_101732_866_B.FIT | 2 | 2 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101734_350_B.FIT | 2 | 1303 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101735_633_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101737_116_B.FIT | 2 | 1 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101738_600_B.FIT | 2 | 2 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101740_083_B.FIT | 2 | 3 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101742_733_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101907_583_B.FIT | 3 | 1 | I | n/a | n/a | n/a |  |  | Dark
02183 | AB | DRKW | I020702_101908_333_B.FIT | 3 | 2 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101909_066_B.FIT | 3 | 3 | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101909_800_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101910_550_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101911_283_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101912_016_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_101914_000_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102043_350_B.FIT | 4 | 2 | I | n/a | n/a | n/a |  |  | Dark
02183 | AB | DRKW | I020702_102046_200_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102049_083_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102052_983_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102203_900_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102206_733_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102209_633_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102213_583_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102342_750_B.FIT | 5 | 3 | I | n/a | n/a | n/a |  |  | Dark
02183 | AB | DRKW | I020702_102345_666_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102348_516_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102351_333_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102354_300_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102357_066_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102400_033_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102402_800_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102531_133_B.FIT | 6 | 4 | I | n/a | n/a | n/a |  |  | Dark
02183 | AB | DRKW | I020702_102534_650_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102538_233_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102541_800_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102545_400_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102548_966_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102552_500_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AB | DRKW | I020702_102555_983_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AC | CHK | Memory check mode.
02183 | AC | CHK | I020702_102846_166_B.FIT | 7 | 5 | I | n/a | n/a | n/a |  |  | Memory check mode.
02183 | AC | CHK | I020702_102849_083_B.FIT | 7 |  | I | [5 empty] | Memory check mode.
02183 | AD | AL | Argon lamp test with safety glass and slit in. | [8 empty] | Argon lamp test with safety glass and slit in.
02183 | AD | AL | I020702_103022_983_B.FIT | 4 | 6 | I | 4 | 1.75 | n/a |  |  | Argon lamp. Good lines. Used dark 102043.
02183 | AD | AL | I020702_103027_233_B.FIT | 4 |  | I
02183 | AD | AL | I020702_103403_066_B.FIT | 4 | 7 | I | 4 | 1.75 | n/a |  |  | Moved Argon lamp. Centered, good lines. Used dark 102043.
02183 | AD | AL | I020702_103407_116_B.FIT | 4 |  | I
02183 | AE | KL | Krypton lamp test with safety glass and slit in. | [8 empty] | Krypton lamp test with safety glass and slit in.
02183 | AE | KL | I020702_103735_333_B.FIT | 4 | 8 | I | 4 | 1.75 | n/a |  |  | Krypton lamp oriented so both sides of lamp
02183 | AE | KL | I020702_103739_483_B.FIT | 4 |  | I
02183 | AF | BBJ | Big JPL BB test. | [8 empty] | Big JPL BB test.
02183 | AF | BBJ | I020702_104520_233_B.FIT | 4 | 9 | I | 4 | 1.75 | n/a |  |  | Big JPL BB set point = 225 deg C.
02183 | AF | BBJ | I020702_104524_483_B.FIT | 4 |  | I
02183 | AF | BBJ | I020702_105912_150_B.FIT | 4 | 10 | I | 4 | 1.75 | n/a |  |  | Big JPL BB aperature open. Temp (BB) = 225 deg C. (-1.666 mV).
02183 | AF | BBJ | I020702_105916_150_B.FIT | 4 |  | I
02183 | AF | BBJ | I020702_110240_883_B.FIT | 4 | 11 | I | 4 | 1.75 | n/a |  |  | Big JPL BB aperature closed (masked with Mylar).
02183 | AF | BBJ | I020702_110245_150_B.FIT | 4 |  | I
02183 | AF | BBJ | I020702_111022_266_B.FIT | 4 | 12 | I | 4 | 1.75 | n/a |  |  | Big JPL BB aperature open. Temp (BB) = 225 deg C. (-1.664 mV).
02183 | AF | BBJ | I020702_111026_316_B.FIT | 4 |  | I
02183 | AG | STIMIR | STIM: Increased signal test. | [8 empty] | STIM: Increased signal test.
02183 | AG | STIMIR | I020702_112057_916_B.FIT | 4 | 13 | I | 4 | 1.75 | n/a |  |  | STIM at hex 80. Used dark 102043.
02183 | AG | STIMIR | I020702_112102_066_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_112248_666_B.FIT | 4 | 14 | I | 4 | 1.75 | n/a |  |  | STIM at hex 90. Used dark 102043.
02183 | AG | STIMIR | I020702_112252_716_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_112435_216_B.FIT | 4 | 15 | I | 4 | 1.75 | n/a |  |  | STIM at hex A0. Used dark 102043.
02183 | AG | STIMIR | I020702_112439_516_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_112859_400_B.FIT | 4 | 16 | I | 4 | 1.75 | n/a |  |  | STIM at hex B0. Used dark 102043.
02183 | AG | STIMIR | I020702_112903_650_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_113301_150_B.FIT | 4 | 17 | I | 4 | 1.75 | n/a |  |  | STIM at hex C0. Used dark 102043. At 3.3 and 3.4 microns (at x = ~772 and ~796 for today), polystyrene absorption has SN ~3. Want SN 10 for successful test. Try another step Note: The 7/1 polystyrene-film test produced absorbtions from x = ~ 791 - 835.
02183 | AG | STIMIR | I020702_113305_250_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_114947_066_B.FIT | 2 | 18 | I | 4 | 1.75 | n/a |  |  | Operator error.
02183 | AG | STIMIR | I020702_114949_166_B.FIT | 2 |  | I
02183 | AG | STIMIR | I020702_115034_866_B.FIT | 4 | 19 | I | 4 | 1.75 | n/a |  |  | STIM at hex D0. Take 8 frames to be coadded later to improve SN. Used dark 102043. STIM still in operation. Decided to stop this STIM experiment at hex D0 instead of going to max hex FF. SN seems to improve by a factor of ~2 with each STIM step.
02183 | AG | STIMIR | I020702_115039_016_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115257_700_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115300_533_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115303_433_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115306_266_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115309_150_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115311_983_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115314_883_B.FIT | 4 |  | I
02183 | AG | STIMIR | I020702_115318_633_B.FIT | 4 |  | I
02183 | AH | TLM | M filter test with Tungsten lamp | [8 empty] | M filter test with Tungsten lamp (to study longer wavelengths). Filter is centered near 4.8 mircon (and is about 50 pixels wide in image).
02183 | AH | TLM | I020702_120927_316_B.FIT | 4 | 1322 | I | 4 | 1.75 | n/a |  |  | M filter test with Tungsten lamp at 12 Amp. Slit width fully open. Used dark 102043.
02183 | AH | TLM | I020702_120931_666_B.FIT | 4 |  | I
02183 | AH | TLM | I020702_121319_133_B.FIT | 4 | 1323 | I | 4 | 1.75 | n/a |  |  | M filter with Tungsten lamp at 12 Amp. Closed slit to 250mm width. Have ~unsaturated emission at x = 970 - 1020: DN ~6000. Used dark 102043.
02183 | AH | TLM | I020702_121323_433_B.FIT | 4 |  | I
02183 | AH | TLM | I020702_121536_266_B.FIT | 4 | 1324 | I | 4 | 1.75 | n/a |  |  | M filter with Tungsten lamp. Turned down lamp intensity to 10 Amp. Slit still at 250mm width. GOOD IMAGE. Have max DN 5000 at x =~985. Used dark 102043.
02183 | AH | TLM | I020702_121540_316_B.FIT | 4 |  | I
02183 | AH | TLM | I020702_121948_683_B.FIT | 4 | 1325 | I | 4 | 1.75 | n/a |  |  | Removed M filter and acquired an image with Tungsten lamp at 10 Amp. Used dark 102043.
02183 | AH | TLM | I020702_121952_933_B.FIT | 4 |  | I
02183 | AI | TLNONL | Non-linearity test | [8 empty] | Non-linearity test with gold IR integrating sphere and Tungsten lamp.
02183 | AI | TLNONL | I020702_123355_466_B.FIT | 4 | 1326 | I | 4 | 1.75 | n/a |  |  | Gold integrating sphere with Tungsten lamp at 15 Amp (and 4.2 V). Have about same signal intensity as 7/1 and x-slice has similar flatness as 7/1. Proceed with the linearity test.
02183 | AI | TLNONL | I020702_123359_516_B.FIT | 4 |  | I
02183 | AJ | TLNONL | Begin non-linearity test
02183 | AJ | TLNONL | I020702_123845_966_B.FIT | 6 | 2 | I | 4 | 1.75 | n/a |  |  | Begin non-linearity test. Gold integrating sphere with Tungsten lamp at 15 Amp and 4.2 V. Took 9 sets of 4 frames at each delay time, enough to give good noise stats and frames at start of various sets. Delay time is incremented 2.048 seconds per set; this gets added to the 2.8 second frame time. Used various darks, from 102531 to 102555.
02183 | AJ | TLNONL | I020702_123849_533_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_123853_016_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_123856_450_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124027_950_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 2
02183 | AJ | TLNONL | I020702_124034_916_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124041_866_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124048_833_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124227_983_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 3
02183 | AJ | TLNONL | I020702_124239_050_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124250_100_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124301_150_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124447_500_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 4
02183 | AJ | TLNONL | I020702_124502_600_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124517_800_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124532_950_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124724_683_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 5. Appear to be saturating, starting with this set??
02183 | AJ | TLNONL | I020702_124743_933_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124803_183_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_124822_433_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125020_083_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 6
02183 | AJ | TLNONL | I020702_125043_416_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125106_816_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125130_166_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125334_066_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 7
02183 | AJ | TLNONL | I020702_125401_566_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125428_966_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125456_450_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125706_566_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 8
02183 | AJ | TLNONL | I020702_125738_100_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125809_683_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_125841_166_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_130057_416_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Set 9
02183 | AJ | TLNONL | I020702_130133_050_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_130208_683_B.FIT | 6 |  | I
02183 | AJ | TLNONL | I020702_130244_266_B.FIT | 6 |  | I
02183 | AK | TLG | Check signal of gold IR sphere. Tungsten lamp at 15 Amp and 4.2 V.
02183 | AK | TLG | I020702_130702_516_B.FIT | 4 | 1400 | I | 4 | 1.75 | n/a |  |  | Check signal of gold IR sphere. Tungsten lamp at 15 Amp and 4.2 V.
02183 | AK | TLG | I020702_130706_950_B.FIT | 4 |  | I
02183 | AL | STRAY | Scattered light test with soldering iron and slit in. | [8 empty] | Scattered light test with soldering iron and slit in.
02183 | AL | STRAY | I020702_132119_733_B.FIT | 4 | 1401 | I | 4 | 1.75 | n/a |  |  | Iron turned off. Use this as the dark.
02183 | AL | STRAY | I020702_132123_883_B.FIT | 4 |  | I
02183 | AL | STRAY | I020702_132552_133_B.FIT | 4 | 1401 | I | 4 | 1.75 | n/a |  |  | Iron turned on at 600 deg F. Position origin = 465, 25 mil. Used dark 132119.
02183 | AL | STRAY | I020702_132556_233_B.FIT | 4 |  | I
02183 | AL | STRAY | I020702_133003_200_B.FIT | 4 | 1402 | I | 4 | 1.75 | n/a |  |  | Move iron to offset (+1, 0) where one unit = 0.25 inch. This is stray light point #2 in Ball Doc 564451. Used dark 132119.
02183 | AL | STRAY | I020702_133007_600_B.FIT | 4 |  | I
02183 | AL | STRAY | I020702_133208_383_B.FIT | 4 | 1403 | I | 4 | 1.75 | n/a |  |  | Move iron to offset (+1, -1). Used dark 132119.
02183 | AL | STRAY | I020702_133212_433_B.FIT | 4 |  | I
02183 | AL | STRAY | I020702_133328_883_B.FIT | 4 | 1404 | I | 4 | 1.75 | n/a |  |  | Move iron to offset (+1, -2). Used dark 132119.
02183 | AL | STRAY | I020702_133333_133_B.FIT | 4 |  | I
02183 | AL | STRAY | I020702_133523_266_B.FIT | 4 | 1405 | I | 4 | 1.75 | n/a |  |  | Move iron to offset (+2, -2). Good CO2 absorbtion at x = 944. No obvious problems with scattered light. Used dark 132119.
02183 | AL | STRAY | I020702_133527_366_B.FIT | 4 |  | I
02183 | AM | BBS | Flat Field test with SIRTF BB. | [8 empty] | Flat Field test with SIRTF BB.
02183 | AM | BBS | I020702_133853_200_B.FIT | 4 | 1407 | I | 4 | 1.75 | n/a |  |  | SIRTF BB against window.
02183 | AM | BBS | I020702_133857_300_B.FIT | 4 |  | I
02183 | AN | DRKW | Allmodes - wood-box darks.
02183 | AN | DRKW | I020702_134159_733_B.FIT | 1 | 1408 | I | 4 | 1.75 | n/a |  |  | Collected wood-box darks.
02183 | AN | DRKW | I020702_134202_500_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134205_466_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134208_250_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134211_200_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134213_966_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134216_933_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134220_816_B.FIT | 1 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134348_783_B.FIT | 2 | n/a | I | 4 | 1.75 | n/a |  |  | Dark
02183 | AN | DRKW | I020702_134350_266_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134351_750_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134353_050_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134354_516_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134356_000_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134357_483_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134400_133_B.FIT | 2 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134526_016_B.FIT | 3 | n/a | I | 4 | 1.75 | n/a |  |  | Dark
02183 | AN | DRKW | I020702_134526_750_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134527_500_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134528_233_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134528_966_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134529_716_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134530_450_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134532_466_B.FIT | 3 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134701_766_B.FIT | 4 | n/a | I | 4 | 1.75 | n/a |  |  | Dark
02183 | AN | DRKW | I020702_134704_616_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134707_500_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134711_350_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134822_316_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134825_166_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134828_050_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_134832_100_B.FIT | 4 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135001_150_B.FIT | 5 | n/a | I | 4 | 1.75 | n/a |  |  | Dark
02183 | AN | DRKW | I020702_135004_116_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135006_883_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135009_800_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135012_616_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135015_433_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135018_383_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135021_200_B.FIT | 5 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135149_466_B.FIT | 6 | n/a | I | 4 | 1.75 | n/a |  |  | Dark
02183 | AN | DRKW | I020702_135153_100_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135156_683_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135200_183_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135203_783_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135207_333_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135210_816_B.FIT | 6 |  | I | [5 empty] | Dark
02183 | AN | DRKW | I020702_135214_216_B.FIT | 6 |  | I | [5 empty] | Dark

## Summaries
Image Log Summary by Day
June 27 2002
02178 |  |  | Operator: | Hampton
02178 |  |  | Optical Bench Temp: | not recorded
02178 |  |  | 564449 - Collected IRFPA Image Table
02178 | AA | LVL | Level
02178 | AB | MSK | Focus
02178 | AC | MSK | Focus
02178 | AD | MSK | Focus
02178 | AE | MSK | Focus
02178 | AF | MSK | Focus
02178 | AG | MSK | Focus
02178 | AH | MSK | Focus
02178 | AI | MSK | Focus
02178 | AJ | MSK | Focus
02178 | AK | MSK | Focus
02178 | AL | MSK | Focus
02178 | AM | MSK | Focus
02178 | AN | MSK | Focus
02178 | AO | MSK | Focus
02178 | AP | MSK | Focus
June 28 2002
02179 |  |  | Operator: | Van Cleve
02179 |  |  | Optical Bench Temp: | not recorded
02179 |  |  | 564449 - Collected IRFPA Image Table
02179 | AA | LVL | Level
02179 | AB | DRKW | Allmode darks
02179 | AC | CHK | Memory Check
02179 | AD | MSK | Focus
02179 | AE | MSK | Focus
02179 | AF | MSK | Focus
02179 | AG | MSK | Focus
02179 | AH | MSK | Focus
02179 | AI | DRKW | Darks
02179 | AJ | MSK | Focus
02179 | AK | MSK | Focus
02179 | AL | MSK | Focus
02179 | AM | MSK | Focus
02179 | AN | MSK | Focus
02179 | AO | DRKW | Darks
02179 | AP | MSK | Focus
02179 | AQ | MSK | Focus
02179 | AR | MSK | Focus
02179 | AS | MSK | Focus
02179 | AT | MSK | Focus
June 29 2002
02180 |  |  | Operator: | Van Cleve
02180 |  |  | Optical Bench Temp: | not recorded
02180 |  |  | 564449 - Collected IRFPA Image Table
02180 | AA | LVL | Level
02180 | AB | DRKW | Allmode dark
02180 | AC | CHK | Memory check
02180 | AD | STIMIR | Stimulator
02180 | AE | MSK | Focus
02180 | AF | MSK | Focus
02180 | AG | MSK | Focus
02180 | AH | DRKW | Darks
02180 | AI | MSK | Focus
02180 | AJ | MSK | Focus
02180 | AK | MSK | Focus
02180 | AL | MSK | Focus
02180 | AM | MSK | Focus
02180 | AN | MSK | Focus
02180 | AO | MSK | Focus
02180 | AP | MSK | Focus
02180 | AQ | MSK | Focus
02180 | AR | MSK | Focus
02180 | AS | MSK | Focus
02180 | AT | MSK | Focus
02180 | AU | MSK | Focus
02180 | AV | STIMIR | Stimulator
02180 | AW | AL | Argon lamp
June 30 2002
02181 |  |  | Operator: | Van Cleve
02181 |  |  | Optical Bench Temp: | 128K
02181 |  |  | 564449 - Collected IRFPA Image Table
02181 | AA | LVL | Levels
02181 | AB | DRKW | Allmodes dark
02181 | AC | CHK | Memory check
02181 | AD | DRKW | Darks
02181 | A | CHK | Memory check
02181 | AF | UNK | Test condition not recorded
02181 | AG | STIMIR | Stimulator
02181 | AH | STIMIR | Stimulator
02181 | AI | UNK | Test condition not recorded
02181 | AJ | AL | Argon
02181 | AJ | AL | Allmodes Argon center slit
02181 | AL | AL | Allmodes Argon top slit
02181 | AM | AL | Argon bottom of slit
02181 | AN | AL | Allmodes Argon
02181 | AO | KL | Krypton
02181 | AP | KL | Allmodes Krypton
02181 | AQ | BBJ | Blackbody (which)
July 1 2002
02182 |  |  | Operator: | Smith (AM) / Van Cleve (PM)
02182 |  |  | Optical Bench Temp: | 135K
02182 |  |  | 564449 - Collected IRFPA Image Table
02182 | AA | BAD | Bad file
02182 | AB | LVL | Level check
02182 | AC | DRKW | Dark. Reference = wood box in corner.
02182 | AD | CHK | Memory check mode.
02182 | AE | MSK | Knife Edge
02182 | AF | MSK | Change Z-Focus to -1mm.
02182 | AG | BBJ | Black Body=2250
02182 | AH | DRKM | BB dark with mylar
02182 | AI | BBJ | BBJ=?
02182 | AJ | BBJ | BB fill tests
02182 | AK | AL | Argon lamp
02182 | AL | KL | Krypton
02182 | AM | TLG | Tungsten with gold IP
02182 | AN | BBJ | BB mica test
02182 | AO | BBJ | BB Polystyrene
02182 | AP | STIMIR | Internal stimulus
02182 | AQ | DRKW | SIRTF BB=300
02182 | AR | BBS | SIRTF BB=300 to 450
02182 | AS | BBS | Sirtf BB with Polystyrene
02182 | AT | TLW | Tungsten lamp, white integrating sphere
02182 | AU | DRKW | Darks
02182 | AV | TL12 | Tungsten lamp, latent light
02182 | AW | TL12 | Tungsten lamp, latent light
02182 | AX | TL12 | Tungsten lamp, latent light
02182 | AY | TL12 | Tungsten lamp, latent light
July 2 2002
02183 |  |  | Operator: | Van Cleve
02183 |  |  | Optical Bench Temp: | 141K
02183 |  |  | 564449 - Collected IRFPA Image Table
02183 | AA | LVL | Level check
02183 | AA | LVL | Level check
02183 | AB | DRKW | Wood-box Darks.
02183 | AC | CHK | Memory check mode.
02183 | AD | AL | Argon lamp test with safety glass and slit in.
02183 | AE | KL | Krypton lamp test with safety glass and slit in.
02183 | AF | BBJ | Big JPL BB test.
02183 | AG | STIMIR | STIM: Increased signal test.
02183 | AH | TLM | M filter test with Tungsten lamp
02183 | AI | TLNONL | Non-linearity test
02183 | AJ | TLNONL | Begin non-linearity test
02183 | AK | TLG | Check signal of gold IR sphere. Tungsten lamp at 15 Amp and 4.2 V.
02183 | AL | STRAY | Scattered light test with soldering iron and slit in.
02183 | AM | BBS | Flat Field test with SIRTF BB.
02183 | AN | DRKW | Allmodes - wood-box darks.

## Experiment
Exp. ID | Experiment Description
BAD | Data is erroneous
DRKW | Dark with wood block
DRKM | Dark with mirror held at chamber window
BBJ | Black Body, JPL
BBS | SIRTF black body.
TL | Tungsten lamp, calibrated at 15.0 amps.
TL12 | Tungsten lamp, calibrated at 12.0 amps.
TLG | Tungsten lamp, gold integrating sphere.
TLW | Tungsten lamp, white integrating sphere.
TLM | M filter with Tungsten lamp.
TLNONL | Non-linearity test with Tungsten lamp.
KL | Krypton lamp.
AL | Argon lamp.
MSK | HRI IR focus or boresight alignment test for thermal-vacuums 1 and 2, during thermal-vacuum 1, using a pinhole or slit of various sizes.
CHK | Memory check.
LVL | Level check.
STRAY | Straylight test with soldering iron.
STIMIR | HRI-IR internal stimulator test.
UNK | Test condition not recorded